## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 97 | Bidder: |  | TOTAL DAP PRICE OF LOT No97: |  |  | 0
Delivery address: Institut za nizijsko šumarstvo i životnu sredinu u Novom Sadu, Antona Čehova 13, 21000 Novi Sad, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | CO2 / H2O |  |  | 1 | 0 | 0
| 443 | Requrements:1. CO2/H2O analyzer a. Technical characteristics i. Type of analyzer: Absolute, non-dispersive infrared gas analyzer.ii. Detector type: Thermoelectrically cooled lead selenide detector.iii. Bandwidth- sampling rate: 5, 10, or 20 Hz, user-selectable.iv. Calibration range: 0-3000 µmol mol-1 for CO2 0- 60 mmol mol-1 for H2O v. Accuracy: Within 1% of reading for CO2 Within 2% of reading for H2O vi. Zero drift (per °C): ±0.1 ppm typical, ±0.3 ppm max.for CO2 ±0.03 ppt typical, ±0.05 ppt max. for H2O vii. Gain drift (% of reading per °C) : ±0.02% typical, ±0.1% max.@ 370 ppm for CO2 ±0.15% typical, ±0.30% max.@ 20 ppt for H2O viii. Direct sensitivity to H2O (mol CO2/mol H2O): ±2.00E-05 typical, ±4.00E-05 max.ix. Direct sensitivity to CO2 (mol H2O/mol CO2): ±0.02 typical, ±0.05 max.x. Operating Temperature Range: -40 to 50°C xi. Power Requirements: 10.5 to 30 volts DC.xii. Low power consumption: up to 30 W 2. Datalogger:a. Data Storage: Removable industrial grade USB flash storage device (4 GB provided, addressable capacity > 16 GB). b. Data Communication: Ethernet, Synchronous Devices for Measurement (SDM; >50 Hz), RS-232 (115,200 baud; 20 records per second max), 6 DACs (0-5 V; 300 Hz). c. Inputs: Ethernet, 4 analog input channels. d. System response: 5, 10, or 20 Hz, user-selectable. e. User Interface: Windows based. 3. Equipment for calibration of analyzer : - Dew point generator or similar equipment for H2O calibration 4. Equipment for connecting with Gill anemometer (cable, converter..) 5. Software for Eddy covariance data processing 6. Customer service support 7. Broad reference list of applied equipment in Eddy covariance research
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details