## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 155 | Bidder: |  | TOTAL DAP PRICE OF LOT No155: |  |  | 0
0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Microscope |  |  | 1 | 0 | 0
| 469 | 1. Microscope stand - lower stand part illumination- transmitted-light illumination with switchable diffusion disk- dovetail mount for attachable stage carriers- 24 mm focus lift- adjustable focus stop- 6-position filter wheel for dia. 36 mm filters- mount for 2-position filter slider 14x40, 36 mm- upper stand part- nosepiece 6x brightfield/darkfield DIC, M27- HD reflected-light illumination- mechanical safety shutter and switchable diffusion disk- dovetail mount for changeable reflector turrets and reflector slider- mount for luminous-field diaphragm slider and aperture stop slider or FL attenuator, discrete and filter slider 14x40 mm and polarizer slider A 6x30 mm 2. Stop slider A 14x40 mm with luminous-field diaphragm usable with reflected-light illuminator 3. Stop slider A 14x40 mm with aperture stop usable with reflected-light illuminator 4. Reflector turret 6x man., changeable, for P&C modules 5. Filter set D containing 3 grey filters 0.015/0.06/0.25, d=36 mm 6. Filter slider D 14x40 mm, 2 positions for filter diameter 36 mm for transmitted light 7. Conversion filter CB 3, d=36x2 8. Binocular phototube 20°/23 (100:0/0:100), upright image with sliding prism, camera port with interface 60N 9. Sideport ISCP 60N P&C left 0.63x for transmitted light Infinity Space Camera Port with 2 switching positions 100% VIS and mount for P&C module, integrated camera factor 0.63x, compatible with camara adapters 0.63x and higher 10. Reflector module brightfield ACR P&C for reflected light 11. Stage carrier D/A; attachable and vertically adjustable 12. Condenser carrier with vertical adjustment on both sides 13. Mechanical stage 75x50 R with hardcoat anodized surface 220 x 170 mm stage plate, covered X-guide positioned at bottom, right drive 135 mm (extendable by 15 mm) with friction setting 14. Specimen holder basic for one-hand operation, spring lever left 15. Dust protection Set M consisting of dust cover (L650xW200xH570 mm), eyepiece cover and protection cap for luminous-field diaphragm optics 16. Illuminator with quartz collector lamp mount and heat-reflecting filter 17. 12 V 100 W halogen lamp 18. Power supply external and LED lamps stabilized, 90...246V/50...60Hz/225VA, 0...12V DC. Incl. 2 lamp sockets. With country-specific line cable. 19. Reflector module Reflection avoided through tilted mount for emission filters, max. filter thickness: 5 mm 20. Filter set (Blue) 21. Filter set (Green) 22. Filter set 20 Rhodamin (Red) 23. Antiglare screen 24. Illuminator inclusive built in power supply, lamp module, infrared filter and CAN-BUS cable. Fluorescence light source with integrated shutter with trigger capability (5V TTL), that works stepper motor controlled, fast and vibration free. With CAN connection for switching the shutter. 25. Lightguide HXP with liquid fill, 2m (D) ø3mm x 2000mm 26. Illumination adapter for light sources with light guide 27. Shutter remote control incl. connecting cable 2 m. Usable as knuckle switch or pedal switch. 28. Objective 5x/0,15Ph1 29. Objective 10x/0,25 M27 (WD=4.5mm) 30. Objective 20x/0.45 M27 (WD=0.46mm) 31. Objective 40x/0.75 M27 (WD=0.71mm), incl. Cover glasses, high performance, D=0.17mm, box with 100 pc. 32. Objective 100x/1.30 Oil M27(WD=0.20mm), incl. Immersol 518 F, oiler 20ml and Cover glasses, high performance,D=0.17mm, box with 100 pc. 33. DIC slider 34. Condenser, achromatic-aplanatic with front lens switchable on the left and right. 5-position modulator disk. 4 centerable positions with Ph1, Ph2, Ph3 and darkfield stop 0.65. 1 brightfield position with aperture stop. For objectives 1.0-100x, WD=1.0mm 35. Condenser modules with polarizer 36. Slider 10x/0.25 37. Slit-diaphragm 2mm for integration into the modulator disk 38. Slider 20x/0.45 39. Slit-diaphragm 4.5mm for integration into the modulator disk 40. Analyzer module for transmitted light optimized for DIC in transmitted light, not suitable for polarization 41. Analyzer module for transmitted light 42. Eyepiece PL 10x/23 Br. foc. 43. Eyepiece eyecup 44. Digital Microscopy Camera 5s with driver software, configuration tool and connection cable USB 2.0, 3m. Sensor: 5 Megapixels, color Basic resolution: 2560 (H) x 1920 (V) = 5.0 Megapixels Pixel size: 2.2 μm x 2.2 μm Sensor size: 5.7 mm x 4.28 mm equivalent 1/2.5" (diagonal 7.1 mm) Live frame rate (depending on hardware and software configuration): H x V Frame Rate 800 x 600 max. 13 Digitization: 3 x 8 bit/pixel Integration time: 10 μs up to 2 s Interfaces: 1x SD card slot, 1x mini USB 2.0, 1x AV (S-Video), 1x DVI- D (HDMI) Spectral range: Approx. 400 nm-700 nm, IR filter Read-out mode: Progressive Optical interface: C-mount Housing: blue anodized aluminium Registration: CE Ambient conditions: +5° ... +45° Celsius, 10% .... 80% relative air humidity, not condensing, free air circulation required Power connection: via USB 2.0 or external power supply (optional) Dual color LED: Power on and ready for capture (green); Recording (blinking green); Not ready (red); Error (blinking red) Integrated slot: SD card slot for SD and SDHC cards Recording: Switch for image capture Set new white balance: Switch for new setup Supported operating systems: Windows XP, Windows 7 45. Camera Adapter 60N-C 2/3" 0.5x 46. Digital Microscopy Camera ICm1 Rev.1 (D) incl. driver software and FireWire / IEEE1394 b/a interface cable Number of Pixels: 1388 (H) x 1038 (V) = 1.4 Mega pixel monochrome Pixel size: 4.65 μm x 4.65 μm Chip size: 1/2" Spectral range: With protection glass approx. 400 to 1000 nm Live frame rates (depending on hardware and software configuration): H x V Frame Rate@1ms 1388 x 1038 15 fps Readout of Sensor Sub-Regions ("ROI"): Adjustable Digitization: 12 Bit / Pixel Dynamic Range: Typical 1000 : 1 (60 dB) Integration Time: 1 ms to approx. 4 s Interface (camera): IEEE1394b interface Interface (PC / board): IEEE1394a interface Optical Interface: C-Mount Size / Weight: approx. 4.5 cm x 4.5 cm x 5.3 cm / 500 g Housing: Blue varnished aluminum Registration: CE Power Supply: via FireWire interface Environmental conditions: +5° ... +45° Celsius 47. Apropriate softwares 48. PC, Windows 7, Monitor 24" Delivery address: Institut za hemiju, tehnologiju i metalurgiju u Beogradu, Belgrade, Studentski trg 12-16, Serbia
2 | Transmission and reflection polarizing microscope |  |  | 1 | 0 | 0
| 871 | Magnification: 40X-1,000X for observation Eyepiece Tube-Trinocular inclined 30′ built in betrland lens with fixed diaphragm, 1X magnification factor Intermediate Tube: Built in analyzer rotatable, removable from optical path vernier reading to 6′ DIN compensator slot Revolving Nosepiece Stage: Centering quadruple nosepiece interchangeable 160mm" circular graduated stage 1o increments and vernier reading to 6′, lick stops at every 45o Illumination: Kochler system; centerable and focusable 12V-50W halogen lamp Eyepiece Lens: CFW 10X with diopter adjustment: i. CFW 10 CM with diopter adjustment with crosslines and micrometer ii. CFW 10XM (photo mask) with diopter adjustment Objective CF achromat P-4X N.A. 0.10 i. CF achromat P-10X N.A. 0.25 ii. CF achromat P-10X N.A. 0.65, spring loaded iii. CF achromat P-100X N.A. 1.25 Oil-immersion, spring-loaded iv. CF achromat P-2.5 N.A. Condenser: Achromat strain-free condenser, N.A. 0.90, with iris diaphragm Photography: Fully automatic via silicon photodiode i. A 30% area of 35 mm picture format is measured Base: Builtin transformer with voltage control, field diaphragm built-in supplied with 45 mm filters NCBIO, GIF, ND2 and ND16 Power Requirements: AC 230 V, S0Hz 1o Camera- Digital Color CCD Sensor Pixel Size 4.65 micron x 4.65 microns (min), resolution 1.4 mega pixel (minimum), Camera Speed - Frame Rate 10 fps (min) at full resolution. Image analysis software for image analysis- Modules for Interactive measurements, digital auto focus, extended annotations and multi view. Computer- Suitable PC, Monitor, printer and UPS Delivery address: Filozofski fakultet u Beogradu, Čika Ljubina 18-20, 11000 Belgrade, Serbia
3 | mikroskop |  |  | 1 | 0 | 0
| 1837 | Research-grade upright microscope with fixed stage for single patch-clamp electrophysiology and cell imaging using ratiometric fluorescence probes on tissue slices Microscope frame (upper and lower part) with corresponding optics to achieve light path for epifluorescence applications, DIC visualization and brightfield imaging of thick specimens such as tissue slices Massive fixed stage, manual microscope translator Lamp house with 100 W halogen lamp for IR applications, with collector lens, IR-cold mirror and bulb Power supply for 100 W Halogen lamp, cable for connecting halogen lamp Reflected light illuminator, 5 or 6 position filter turret, aperture, analyzer slider for reflected light polariser for visible and IR-DIC Bi- or trinocular phototube Focusable eyepiece, nosepiece with DIC prism 1 empty filter holder/cube Long working distance condenser for BF/POL/DIC C-mount adapter for camera together with a ccd camera adapter (to work with the chosen illumination/imaging system optics from proj ON173040 code 1838) Fluorescence optical epi fluorescence attachment (to work with illumination/imaging system optics from proj ON173040 code 1838) Narrow IR interference filter for IR microscopy, DIC prisms for 10x and 40x objectives Image video adapter tube Objective, 10x, water immersion long working distance, for electrophysiology applications Objective, 40x water immersion, for electrophysiology applications with IR, VIS and UV fluorescent transmission. Following characteristics are required: N.A. 0.8 or greater long working distance 3 mm excellent spectral transmission in UV, (at 350nm and at 380nm), in order to work in fura2 and SBFI applications as well as in VIS and DIC >50% angle of pipette approach minimal 30 deg Micromanipulator for patch-clamp, right-handed. Following characteristics are required: true approach axis resolution of fine movement 0.1 microm or less length of fine travel 150 microm or more Mount to be assembled with micromanipultor: steinless steel (or aluminium), manipulator/chamber platform, and low profile adapters for shallow angle approach Delivery address: Institut za biološka istraživanja `Siniša Stanković` u Beogradu, bul. Despota Stefana 142, Belgrade, Serbia
4 | Photonic system for cell imaging |  |  | 1 | 0 | 0
| 1838 | Monochromator-based microscope illumination complete imaging platform, for high speed, high sensitivity microscope imaging during patch clamping 1. Monochromator of following characteristics: bandwidth resolution (nm) 1nm wavelength range (nm) 330-650 200nm transition <2ms 2. Xenon lamp with highly stabilized power supply 3. Quartz fiber or UV transmitting liquid light guide to couple monochromator illumination to the microscope condenser 4. Microscope coupling to the microscope to be purchased for the project III41014 code 1837 5. Fura-2 filter set (dichroic and wide band emitter) dimensions for microscope for the project III41014 code 1837 6. Synchronisation electronics between the illuminator and the camera 7. High speed and high sensitivity digital monochrome interline CCD camera, with capability of vibration free operation while cooled (through remote fan option or fan switch-off option) pixel size 6.45x6.45μm image area 9x7 mm chip resolution 1392x1040 Quantum efficiency (at 500nm) 60% or greater digitization 12 bit or more speed (for full frame) 10fps while running in vibration free mode cooled at minimally (°C) -30 read noise at 10 MHz (e-) <6 dark current (at min temperature) (e-/pixel/s) 0.001 maximal readout rate (MHz) 20 8. Software for high speed on line image acquisition, real time system control option: the offer obtained through public procurement process that comprizes the complete system for illumination and imaging, imaging workstation and comprehensive software package will be favored Delivery address: Institut za multidisciplinarna istraživanja u Beogradu, Bulevar Despota Stefana 142, Belgrade, Serbia
5 | Microscope |  |  | 1 | 0 | 0
| 3421 | Binocular phototube with eyepiece 10x and with min. FOW 23 Halogen illuminator min.12V 100W with possibility to adjust on stand or/and on macro/micro wheel Objectives: 10x/0.30, 20x/0.50, 40x/0,75 and 100x/1.30Oil Motorized Z stack with step size min.25nm 6 position coded nosepiece Light and contrast manager TFT display for controlling and configuring microscope without PC 6 position motorized reflector turret Integrated filters for transmitted light min. 8 pieces Achromatic-aplanatic universal condenser 0.9 H D Ph DIC, mot with minimum 8 position and for objectives from 1 up to 100x Bright field, dakfield, phase contrast and DIC contrast on all objectives Digital color microscope camera: 5 Mpix, Interfaces: 1x SD card slot, 1x mini USB 2.0, 1x AV (S-Video), 1x DVI- D (HDMI) Image processing software PC, Windows 7, monitor (24”) Delivery address: Institut za kukuruz `Zemun Polje` u Beogradu, Slobodana Bajica1 , Zemun Polje, Belgrade, Serbia
6 | Digital confocal microscope |  |  | 1 | 0 | 0
| 4427 | The confocal laser microscope is designed for nanometer level imaging, 3D measurement and roughness measurement of metal materials. Magnification ranges from 108x - 17,280x. Technical description - Specifications Optical System: UIS2 Optical System (Infinity-corrected) Total Magnification: 100x - 17000x Field of View: 2560x2560 - 16x16 µm Display Resolution: 1 nm Microscope Frame Observation Method: BF/DIC/Laser/Laser Confocal DIC Laser: 405 nm Semiconductor Laser Detector: Photomultiplier White Light: White LED Detector: 1/1.8-inch 2-megapixel Single-panel CCD Focus Z Stage: Stroke: 100 mm, Maximum Height of Specimen: 100 mm Z Revolving Nosepiece: Stroke: 10 mm, Resolution: 0.01 µm, Repeatability50x: σn-1=0.012 µm Accuracy 0.2+L/100 μm or Less (L=Measuring Length μm) Motorized 6 position nosepiece Frame grabber Differential Interference Contrast Unit Differential Interference Contrast Slider: U-DICR,Polarizing Plate Unit Built-in Objective Lens: Plan Semi-apochromat: 5x, 10x; Plan Apochromat: 20x, 50x, 100x Optical Zoom: 1x - 8x Stage: Motorized Stage ultrasonic type: 100x100 mm Workstation PC, monitor and corresponding software Installation, Travel Expense, Training Delivery address: IMS - Institut za ispitivanje materijala Srbije a.d. U Beogradu, Bulevar vojvode Misica 43, 11000 Beograd, Serbia
7 | mikroskop |  |  | 1 | 0 | 0
| 5564 | MIKROSKOP (BF) + Fluo 6-fold Nosepiece Mech. stage with rotation (for 1 specimen)slideholder 1 specimen Documentation tube BDT 25+ V100/50/0 Tube top (for 1 camera) Eyepieces HC PLAN S 10x/25 Br. M PL FLUOTAR objectives (2.5x, 5x, 10x, 20x, 40x,) (Solid State Fluorescence Light Source) LAS application software for PC 1 LAS Analysis Bundle(contains: Measurement,Ext.Annotation,Analysis) 0.55x c-mount adapter, HC microscope base stand, consisting of stand base including transmitted light axis with motorized aperture and field diaphragms,motorized shutter, automatic Constant Color Intensity Control (CCIC), with 7 keys for automatic Illumination Management (controlling aperture, field diaphragms and light intensity), with 6 programmable functions keys for automatic Contrast Management, with new LCD-Display. Manual Z drive,with one focussing knob for coarse and fine movement and one focussing knob for fine movement only. Coarse drive 3.5mm per rotation, fine drive 0.2mm per rotation. Complete moving range is 45 up to 50mm (moving range of the focus drive is 25mm) depending on condenser type. Power supply 12V/100W, USB 2.0 cable for PC-connection, mains connection 90-250V, 50-60Hz. Including user manual, Application Suite software CD,sheet with labels for programmable keys. Ground plate without filter magazine New Stand top with coded 6-fold objective nosepiece M25 (25mm thread), including new 1 " fluorescence axis with motorized FIM (Fluorescence Intensity Manager) and motorized field diaphragm disc, with motorized 5-fold turret for fluorescence filter cubes and incident light reflectors BF, ICR, POL/IGS (DF and Smith cannot be adapted) and analyzer block. Including cover for objective turret. Cover for all kinds of stand tops.700 Condenser BF, motorized condenser with motorized top lens, automatic Koehler illumination, for objectives 1.25x-100x, for BF and POL contrast Condenser head 0.90 S1 mechanical stage for biological and industrial applications, with ceramic coated plate, 110° rotation, travel range 76x25mm, right hand handling, with bracket and condenser holder Object holder for biological applications, for use with mechanical stages, for 1 specimen, single-hand-operation, right-handed mechanical documentation tube, with fixed viewing angle 30°, field of view 25, variable beam splitters, 3 beam splitter positions: 100% eyepieces, 50% eyepieces:50% documentation port, 100% documentation port Lamphousing 107/2 , with lamp mount for halogen lamp 12V 100W non adjustable, with 1-lens collector and heat protection filter, mains cable 0.55m (Solid State Fluorescence Light Source)Basic illumination device for Fluorescence microscopy with slots for up to 5 different LED Modules. No LED Module included (please order from separate list). Separate Touch Control Panel needed. Power supply included. For use with incubators please order additional adapter. Tube adapter for mounting on documentation tubes , with port for 1 documentation devicet Eyepiece HC PLAN s 10x/25 Br. M Objective 5x/0.15 Free working distance: 12.0 mm For use with and without Cover Glass Suitable for DIC (prism C,C1,C2) Objective 10x/0.3; free working distance: 11.0 mm For use with and without Cover Glass Suitable for DIC (prism D1/D) Objective 20x/0.50 Free working distance: 1.15 mm For use with a 0.17 mm Cover Glass (DIN/ISO) Not suitable for Incident Light, except Fluorescence Objective 40x/0.75 Free working distance: 0.40 mm For use with a 0.17 mm Cover Glass (DIN/ISO) Not suitable for Incident Light, except Fluorescence Objective 2.5x/0.07 Free working distance: 11.3 mm For use with and without Cover Glass LED Module . Excitation wavelength 470nm.Filtercube 470nm LED Ex:et470/40, Em:et525/50, Di:495 Contains:Image Analysis Module Interactive Measurement Module Annotation Module digital camera with sw kit, colour digital camera for on-screen microscope image display, resolution 2048x1536 pixels, 3.1 Mpixels, 1/2" CMOS sensor progressive scan, pixel size 3.2 μm x 3.2 μm, colour filter RGB (Bayer), colour depth 30 bits, exposure time 0.1msec- 2 sec, A/D converter 10 bit, dynamic range >55dB, full size live image, max.18 frames per sec. (depending on the PC performance), gain 1x - 4x, shading correction, optical C- mount, recommended video adapter 0.5x, supported operating systems WindowsXP or WindowsVista, Firewire 1394b interface, Firewire a-b cable 3m, C-mount adapter 0.55x HC, for 1/2" and 0.55" cameras Delivery address: Medicinski fakultet u Beogradu, Dr Subotića 4/2, 11000 Beograd, Serbia
8 | Optical Sensor |  |  | 1 | 0 | 0
| 5628 | Technical description : - CRITICAL 1: Can be installed on and be compatible with CONTURA RDS system - CRITICAL 2: CNC scanner change in conjunction with probe magazine and controller software - Line width: minimum 18mm, the bigger the better - FOV depth: minimum 15mm, the bigger the better - Resolution: maximum 60 μm, the lower the better - Measuring speed: 75,000 pts/s, the higher the better - Typical probing error (MPEp)*: 20 μm, the lower the better - Multi-stylus test (MPEal)**: 20 μm, the lower the better * Comparable to EN/ISO10360-2 ** Comparable to EN/ISO10360-5, for CMM with accuracy of 2μm + L/350 or better Delivery address: Fakultet tehničkih nauka u Novom Sadu, Vladimira Perića Valtera 2, Novi Sad, Serbia
9 | Automatic microscope system |  |  | 1 | 0 | 0
| 9299 | Microscope description:Upright research microscope consisting of stand base including transmitted light axis with motorized aperture and field diaphragms, motorized shutter, automatic Constant Color Intensity Control (CCIC), with 7 keys for automatic Illumination Management (controlling aperture, field diaphragms and light intensity), with 6 programmable functions keys for automatic Contrast Management, with new LCD-Display, with filtermagazine for Transmitted Light. Manual Z drive, with one focussing knob for coarse and fine movement and one focussing knob for fine movement only. Coarse drive 3.5 mm per rotation, fine drive 0.2 mm per rotation. Complete moving range is 45 up to 50 mm (moving range of the focus drive is 25 mm) depending on condenser type. Power supply 12V/100W, USB 2.0 cable for PC-connection, mains connection 90-250V, 50-60Hz. Coded 6-fold objective nosepiece M25 (25 mm thread), including new 1 " fluorescence axis with motorized FIM (Fluorescence Intensity Manager) and motorized field diaphragm disc, with motorized 5-fold turret for fluorescence filter cubes and incident light reflectors BF, ICR, POL/IGS and analyzer block.Motorized condenser with motorized top lens, automatic Koehler illumination, for objectives 1.25x-100x, for BF and POL contrast. Mechanical stage for biological and industrial applications, with ceramic coated plate, 110° rotation, travel range 76x25 mm, right hand handling, with bracket and condenser holder.Object holder for biological applications, for use with mechanical stages, for 1 specimen, single-hand-operation, right-handed.Documentation tube, mechanical documentation tube, with fixed viewing angle 30°, field of view 25, variable beam splitters, 3 beam splitter positions: 100% eyepieces, 50% eyepieces:50% documentation port, 100% documentation port.Lamphousing with lamp mount for halogen lamp 12V 100W non adjustable, with 1-lens collector and heat protection filter, mains cable 0.55 m. Halogen lamp.Objectives:HC PL FLUOTAR 10x/0.30;HC PL FLUOTAR 20x/0.50;HCX PL FLUOTAR 40x/0.75;HCX PL FL 63x/0.70;HCX PL FLUOTAR 100x/1.30 OIL.Eyepieces HC PLAN s 10x/25. Alignment-free external light source with attenuator and fast shutter (lifetime of bulb in average 2000h).Liquid Light Guide.Fluorescence filters:Excitation filter BP 355 – 425, dichromatic mirror 455, suppression filter LP 470;Excitation filter BP 500/20, dichromatic mirror LP515, suppression filter BP 535/30;Excitation filter BP 546/12, dichromatic mirror LP560, suppression filter BP 585/40;Excitation filter BP 436/20, dichromatic mirror LP455, suppression filter BP 480/30.Software description:Cytogenetic network based image analysis and case management software for brightfield and fluorescent samples.Capturing high quality images and On-screen analysis.Control of microscope filter turret and all other motorized microscope modules.Modules: Karyotyping, FISH, Tissue FISH, Karyotype CGH, M-FISH, Flexible Karyotyping, Automated and manual spot counting for cellular FISH.Internal domain server SQL CaseBase Data archive up to 0.5TB RAID1 Remote access for analysis and review Dynamic Z-stack multi-plane imaging captures images of counter stains and probes through each cell Virtual microscopy enables dynamic focusing, up and down through the image.Proprietary Adaptive Image Processing algorithms are used to classify cells based on user defined morphology.Analysis of cells automatically tracked. Digital record of count and results kept for security. Progress tracked during analysis as each class is marked off.Image processing includes background subtraction to eliminate uneven fluorescence in the area, and also ignores signals of extreme size. Archive for results and generating of reports . For FISH Manual Region Drawing with automatic score calculation.Single region and combined slide average ratio score display.Dual Probe Enumeration ratio analysis. Dual Probe Amplification ratio analysis.Cell (DAPI) Classifiers : Default for typical interphase populations; Round-Cell for targeted scoring of circular cells; Lobed-Cell for targeted scoring of lobed nuclei cells;Manual and automatic Karyotyping;Chromosome Enhancements - Contrast & Sharpen;Segmentation tools for Karyotyping: split, overlap and join in automatic or manual mode; USB Dongle protected;PC Dell Precision T3500;24” Monitor;Supports Windows XP SP3 or Windows 7 SP1 Ultimate, 32 bit Delivery address: Biološki fakultet u Beogradu, Takovska 43, Belgrade, Serbia
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F9: Insert CIF price
Note on D10: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F11: Insert CIF price
Note on D12: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F13: Insert CIF price
Note on D14: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F15: Insert CIF price
Note on D16: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F17: Insert CIF price
Note on D18: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F19: Insert CIF price
Note on D20: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F21: Insert CIF price
Note on D22: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F23: Insert CIF price
Note on D24: FILL all white fields, write brand, trade name, model and all relevant trade details