## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 157 | Bidder: |  | TOTAL DAP PRICE OF LOT No157: |  |  | 0
0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Two stereomicroscope with additional equipment |  |  | 1 | 0 | 0
| 442 | Research motorized stereomicroscope with zoom range more than 20:1 Binocular photo tube Eyepiece 10x with field of view minimum 23 Large , stabile stand with base minimum 400x300mm and with motorized focus Gliding stage with XY movement Coded nosepiece with minimum 3 positions Plan APO objectives 0.63x, 1.0x and 1.5x Transmitted light : cold light source LED (900lm light flux, ca. 6500K color temperature) and with: brightfield, darkfield and variable oblique light LED Ring segmented lights for all objectives Double Goose neck light guide powered by cold light source LED (900lm light flux, ca. 6500K color temperature) Drawing attachment for stereoscopic viewing on specimen or for binocular 2D viewing specimen and drawing Multifunctional controller for complete steromicroscope ( zoom, focus, position….), and also for LED cold light sources. Touch screen, joystick, 2 scroll wheels and 6 freely configurable keys. Two foot pedal for additional stereomicroscope control Digital microscope camera with min. 5 megapixel, pixel size with max. 2,2x 2,2µm, interface: 1x SD card slot, 1x mini USB 2.0, 1x AV (S-Video),1x DVI- D (HDMI) Software for camera and microscope control and additional modules for: measurement, Z-stack, multichannel, and extended focus Appropiert PC , Windows 7 and monitor 24” Delivery address: Prirodnomatematički fakultet u Novom Sadu, Trg Dositeja Obradovića 2, 21000 Novi Sad, Serbia
2 | Modular high-performance stereomicroscope |  |  | 1 | 0 | 0
| 873 | Zoom Ratio: Motorized Zoom Ratio of 20:1 or above with Apochromatic Stereo Telescopic optics system and built in aperture diaphragm for increased depth range. It should have a controller with joy stick, key and touch screen for controlling the microscope functions like motorized zooming, motorized focus and control of intensity of cold light source Zoom Range: 0.75x - 15x or above. Magnification: 4.5x - 95x using 0.63x and 10x eyepiece. Objective: Plan apochromat objective 0.63x/0.65x with working distance of 80 or higher. Plan Achromat objective 1.0x with working distance of 80 or higher to be quoted optional. Eyepiece: Wild field paired eyepiece of 10x with F.No. 23mm and diopter settings +/- 5. Nosepiece: Triple revolving coded nosepiece. Stand: Stand with motorized focus with stepsize 350mm and different selectable focus speed including transmitted light illumination for Bright field, Dark field and variable angle oblique illumination with cold light source of 24V 250W. Reflected Light source: Flicker free LED ring light illumination with controller for light intensity and also thermo controller for shadow free illumination. Light segment control and rotation of illumination should be possible for obtaining variable oblique illumination. Storage of 4 illumination setting should be possible. Camera: Digital microscopy CCD camera of 3 megapixel with necessary software for image acquisition. Software: should control functions of microscope like motorized zooming, motorized focus, illumination intensity etc., Image annotations (text, arrows, boxes, circles, scale bars. Image tags such as acquisition and exposure time. Interactive measurement of intensity profiles, length, area, perimeter, circle, angle, counting and marking of events (measurement tools should be freely configured). Image Gallery View. Image splitter display for synchronous or independent comparision of upto 8 2D or multidimensional images with the ability to create publication quality artwork. Generating reports containing images, data tables and graphs. Upgradeability: It should be upgradable for Fluorescence attachment with 4-5 filter cubes. Software should be open for further addition of modules as per application. Computer: A suitable computer that would support the software and hardware functions with sufficient memory, UPS and Color Printer. Delivery address: Filozofski fakultet u Beogradu, Čika Ljubina 18-20, 11000 Belgrade, Serbia
3 | stereomicroscope |  |  | 1 | 0 | 0
| 1595 | Research motorized stereomicroscope with zoom range more than 20:1 Binocular ergo photo tube ( variable angle) Eyepiece 10x with field of view minimum 23 Large , stabile stand with base minimum 300x300mm and with motorized focus Mechanical stage with XY movement Coded nosepiece with minimum 3 positions Plan APO objectives 0.63x, 1.0x and 2,3x Transmitted LED light unit: brightfield, darkfield and variable oblique light LED Ring segmented lights objectives 0,63x and 1.0x Two double Goose neck light guide powered by cold light source LED (900lm light flux, ca. 6500K color temperature) Polarization equipment for ring lights and goose neck light Drawing attachment for stereoscopic viewing on specimen or for binocular 2D viewing specimen and drawing Multifunctional controller for complete steromicroscope ( zoom, focus, position….), and also for LED cold light sources. Touch screen, joystick, 2 scroll wheels and 6 freely configurable keys. Two foot pedal for additional stereomicroscope control Digital microscope camera with min. 5 megapixel, pixel size with max. 2,2x 2,2µm, interface: 1x SD card slot, 1x mini USB 2.0, 1x AV (S-Video),1x DVI- D (HDMI) Software for camera and microscope control and additional modules for: measurement, Z-stack, multichannel, and extended focus Appropiert PC , Windows 7 and monitor 24” Delivery address: Prirodnomatematički fakultet u Novom Sadu, Trg Dositeja Obradovića 2, 21000 Novi Sad, Serbia
4 | Stereomikroscop |  |  | 1 | 0 | 0
| 9294 | Research motorized stereomicroscope with zoom range more than 20:1 Binocular photo tube Eyepiece 10x with field of view minimum 23 Large , stabile stand with base minimum 300x300mm and with motorized focus Mechanical stage with XY movement Coded nosepiece with minimum 3 positions Plan APO objectives 1.0x and 1.5x Transmitted LED light unit: brightfield, darkfield and variable oblique light LED Ring segmented lights Double Goose neck light guide powered by cold light source LED (900lm light flux, ca. 6500K color temperature) LED fluorescence for GFP (green fluorescent protein, extinction 470/40, beamsplitter FT 490nm and extinction 525/50), possibility for upgrade with additional one LED fluorescence ( DAPI, FITC) Multifunctional controller for complete steromicroscope ( zoom, focus, position….), and also for LED cold light sources. Touch screen, joystick, 2 scroll wheels and 6 freely configurable keys. Digital microscope camera with min. 5 megapixel, pixel size with max. 2,2x 2,2µm, interface: 1x SD card slot, 1x mini USB 2.0, 1x AV (S-Video),1x DVI- D (HDMI) Software for camera and microscope control and additional modules for: measurement, Z-stack, multichannel, and extended focus Appropiert PC , Windows 7 and monitor 24” Delivery address: Prirodnomatematički fakultet u Novom Sadu, Trg Dositeja Obradovića 2, 21000 Novi Sad, Serbia
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F9: Insert CIF price
Note on D10: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F11: Insert CIF price
Note on D12: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F13: Insert CIF price
Note on D14: FILL all white fields, write brand, trade name, model and all relevant trade details