## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 56 | Bidder: |  | TOTAL DAP PRICE OF LOT No56: |  |  | 0
Delivery address: Medicinski fakultet u Beogradu, Dr Subotica 6, Belgrade, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | System for eye motion |  |  | 1 | 0 | 0
| 2515 | Softverske karakteristike sistema: Kamera za registrovanje očnih pokreta, sa infracrvenim zracima za fiksiranje zenice i frekvencom snimanja od najmanje 300 Hz; Mogućnost registrovanja različitih paradigmi očnih pokreta: pokreti praćenja, pokreti voljnog usmeravanja (sakadični pokreti), pokreti suprotnog voljnog usmeravanja (anti-sakadični pokreti), registrovanje nistagmusa; Mogućnost registrovanja latence, dužine trajanja, oblika pokreta praćenja i pokreta usmeravanja; Opseg preciznosti registrovanja sakadičnih pokreta i pokreta praćenja od 0.5⁰ horizontalno i 0.5⁰ vertikalno; Ugao registrovanja vertikalnih sakada: ≥ 20⁰; Ugao registrovanja gledanja horizontalnih sakada: ≥ 120⁰ ili beskonačan; Latenca registrovanja sakadičnih pokreta: do ± 2 ms; Sistem za stimulaciju horizontalnih sakadičnih pokreta u bar 2 različite amplitude; Sistem za stimulaciju horizontalnih pokreta praćenja u bar 2 različite amplitude; Postojanje unapred definisanih kliničkih paradigmi ispitivanja sa mogućnošću automatizovane analize; Postojanje unapred definisanih normi amplitude i latence horiznontalnih i vertikalnih pokreta kod zdravih osoba različite životne dobi; Mogućnost definisanja novih paradigmi očnih pokreta unutar softvera; Mogućnost registrovanja pokreta nezavisno na oba oka; Mogućnost izračuvananja odnosa abdukcije i adukcije; Mogućnost merenja veličine zenice; Frekvenca binokularnog registrovanja očnih pokreta od: 120 ili 300 Hz; Rezolucija video registrovanja: o U slučaju opcije od 120 Hz: 800x224 px,o U slučaju opcije od 300 Hz: 400x112 px; Mogućnost sinhronizaciju sa softverskim sistemima kakvi su na primer elektorencefalografija i transkranijalna magnetna stimulacija; Mogućnost manualnu analizu dobijenih podataka;Mogućnost prenosa dobijenih podataka u softverske sisteme za statisticku analizu (Matlab ili Excel).Hardverske karakteristike sistema Resolucija monitora za stimulaciji očnih pokreta od najmanje 1920x1080 pixela; Oblik monitora za stimulaciju očnih pokreta: 16:9 (aspec ratio); Operativni sistem: Windows 7.
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details