## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 3 | Bidder: |  | TOTAL DAP PRICE OF LOT No3: |  |  | 0
0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | CHNS analyzer |  |  | 1 | 0 | 0
| 610 | The instrument should be able to perform the analysis of soil, sediments, and organic materials The instrument should have the capability to run using either Helium or Argon as the carrier gas and Oxygen as Combustion gas Imported regulator for carrier gas Furnace temperature at Combustion zone: 100 - 1800°C Accuracy: < 0.3 % Sample Size: Typically 1-5 mg; up to 500 mg for samples with low carbon content Analysis Time: 5 - 10 minutes for CHN Detection range: C: 0-40mg (or 100%) N: 0-15 mg (or 100%) H: 0-3 mg (or 100%) S: 0-6 mg (or 100%) Standard Deviation :≤0.1% of absolute Simultaneous CHNS analyzer with one sample intake Controller: Fully automated PC controlled; Software should be capable of full instrument control, data acquisition, processing, reprocessing, reporting Automatic leak detection using the software Auto sampler Kit for 2000 analyses CHNS / NCS / S (with Prepacked Quartz reactor) Precision ultra micro balance: Sartorius CPA2P, Mettler or equivalent Bottled Gases Required: Helium/Oxygen 99.995% purity Delivery address:Poljoprivredni fakultet u Beogradu, Nemanjina 6, 11080 Zemun
2 | CHNS Elemental Analyzer |  |  | 1 | 0 | 0
| 1782 | Quantitative sample digestion at up to 1200°C (1800°C at the point of combustion when tin foil is used). Furnace warranty of 10 years. Injection of oxygen directly to the sample. Combustion system that not require special catalysts. Easy ash removal from the combustion system. Multiple analysis modes. Auto-sampler with up to 80 positions. Sample weighing range from 1 mg to 800 mg, up to 20 mg organic matter. Range of element concentrations: 0-7 mg C, 0-2 mg N or S; high C:N ratio of up to 5000:1. Separation of gaseous components with one TPD (Temperature Programmed Desorption). Automated optimization of the time per analysis. Matrix independent and long lifetime multipoint calibration. Helium or Argon carrier gas. Instrument control with a standard PC. Power supply: 230 V, 50-60 Hz Maximal power of up to 2.0 kW. Maximal weight of up to 65 kg. Free kit (including tin/aluminum capsules) for at least 1000 analyses. Sample former with pressing tool. Delivery address:Institut za multidisciplinarna istraživanja u Beogradu, Kneza Viseslava 1, 11030 Beograd
3 | CHNS analyzer |  |  | 1 | 0 | 0
| 8064 | Fully Automated PC Controlled CHNS Elemental Analyzer should be able to analyze soil, sediments, organic materials and fertilizers. 1. Operating modes for measurement: CHNS, CHN, CNS, CN, N, S. 2. Simultaneous C.H.N.S analysis with one sample intake. 3. Analyzer must must be able to analyze true macro samples up to 200 mg organic substance or 1.5 g soil. 4. Upper limits of detection of absolute element contents must be C: 150 mg N: 100 mg H: 15 mg S: 18 mg 5. Gases must be separated by purge & trap technology by independent purge & trap columns for CO2, H2O and SO2, while N2 goes directly to detector; detector is TCD - Thermal Conductivity Detector.6. Mass flow controller positioned before detector to ensure stable flow of resulting gases 7. Auto sampler for 60 macro or 120 small samples. 8. Possibility of use of Argon as carrier gas for cost saving and availability reasons.9. Software should be windows based and should have display actual pressures, flow rates, temperatures, number of samples analyzed with the provision for setting maintenance interval with monitoring of maintenance cycle.10. Microbalance with hands free operation and interface to analyzer for direct transmission of weighing data should be quoted.Specifications for Microbalance: Microbalance with interface to analyzer for direct transmission of weighing data. Balance operation through bright touch screen.Hands‐free infrared controlled opening and closing of draft shield.Weighing range: upto 6100 mg ;Resolution: 0.001 mg, Taring Range : 6100 mg ; Repeatability : 0.8 – 0.9 micro gm, Linearity : ±2‐4 μg.11. System should be supplied complete with Helium and Oxygen gas cylinders and regulators.12. The Gases should be of high purity (oxygen‐ 99.9995 % and he‐ 99.999 %).13. Computer and Printer should also be offered for running the equipment. Specifications for Computer and Printer : Intel PentiumCore2 dual with 2 GB RAM, 500 GB HDD, DVD Drive , 18” TFT monitor.,optical mouse, key board and Laser printer.14. Required spares kit for 3 years trouble free operation to be quoted.15. Voltage stabilizer should be quoted.16. Vendor should clearly mention the after sales local service and parts availability and also list of other users to be provided.17. At least one year warranty period from the date of installation should be quoted. Delivery address: Poljoprivredni fakultet u Novom Sadu, Trg Dositeja Obrdovića8, 21000 Novi Sad Serbia
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F9: Insert CIF price
Note on D10: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F11: Insert CIF price
Note on D12: FILL all white fields, write brand, trade name, model and all relevant trade details