## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 9 | Bidder: |  | TOTAL DAP PRICE OF LOT No9: |  |  | 0
Delivery address: Tehnološko-metalurški fakultet u Beogradu, Faculty of Technology and Metallurgy, Karnegijeva 4, 11120 Belgrade |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Ion trap mass detector, model: LCQ Fleet Ion Trap |  |  | 1 | 0 | 0
| 1333 | Mass detector specification Analyzer • Linear Ion Trap 2D-Linear quadruple Ion trap Mass range • 15-4000 m/z Resolution • Down to 0.05 FWHM Scan modes • Selected Reaction Monitoring (SRM) • MSn from n=od 1 to 15 • Full-scan • Ultra Zoom Scan • automatic MSn – data dependent Ion source • API-ESI Sensitivity • ESI- for 1 pg/μL reserpine minimum S/N 400:1 • APCI- for 1 pg/μL reserpine minimum S/N odnos 400:1 Detection system • double “conversion dynode detector” Options • Integrated Divert Valve; Integrated Syringe Pump Vacuum system • Differentially-pumped vacuum system to 10-5Torr • Split-flow turbo molecular pump controlling vacuum in three regions • Dual rotary vacuum pump configuration PC with monitor • Capable of supporting adequate software and data storage requirements Mass spectrometer must have possibility to connect to the existing HPLC system and existing software solution. Installation and basic training should be included. Linear Ion Trap Mass Spectrometry Detector is planned as an improvement of the existing HPLC-MS system, supplied by Thermo Scientific. Therefore, the upgrade detector should be purchased from the same manufacturer.
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details