## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 70 | Bidder: |  | TOTAL DAP PRICE OF LOT No70: |  |  | 0
Delivery address: Elektronski fakultet u Nišu, Aleksandra Medvedeva 1, 18000 Niš, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | X-Ray Flat Panel Detector |  |  | 1 | 0 | 0
| 688 | Specifications: (1) Image Format: Direct X-ray Conversion , Active Area: min 340(H)×420(V)mm (14 x 17 inch), Pixel Matrix: min 2300(H)×3000(V), Pixel Pitch: max 143 μ m (2) Performance: Limiting Resolution: 3.5 lp/mm typ, DQE (0), Quantum - Limited: > 60 %, A/D Conversion: min 14 bit (3) Ratings: Energy Range: 40 - 150 kVp, Maximum Entrance Dose (low Gain): 4 mR / frame (4) Interface: Data Output: high speed Ethernet, Command Control: high speed Ethernet, X-ray Synchronization Control: External (5) Image Acquisition Mode Table: Frame rate (Frame/s): 1(approx), Preview time: max 5 sec, X-ray period (ms): from 500 ms. (6) General: Acquisition to full resolution image display in under 10 seconds. Selectable advanced image enhancement (AIE) for optimal image brightness, contrast .and edge enhancement. Auto Region Of Interest (ROI) for optimal display of selected image area. Expanded edge enhancement. Measurement and annotation tools. Zoom/pan. Image rotation. (7) CE compatible (preferable)
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details