## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 95 | Bidder: |  | TOTAL DAP PRICE OF LOT No95: |  |  | 0
Delivery address: Medicinski fakultet u Beogradu, Koste Todorovica 6, Belgrade, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Video colonoscope |  |  | 1 | 0 | 0
| 31 | Video colonoscope:• Field of view 170° • Depth of view 2-100 mm• Insertion tube outer diameter 12.8 mm• Instrument channel inner diameter 3.7 mm• Working length min. 1680 mm• Bending up/down 180°/180°, left/right 160°/160°• HDTV video colonoscope• Narrow Band light and imaging for capillary pattern visualization • Auxiliary water jet• Variable stiffness capability for adjustment of insertion tube• Scope ID function• Three light guides for better visualization • Standard set included with colonscope: Biopsy forceps Biopsy valves Channel cleaning brush (1 pc.) Cleaning adapters and system for cleaning and disinfection Light source:• 300 W Xenon lamp • Approx. 500 hours lamp life• Emergency halogen lamp• Automatic and manual adjust of light intensity • Air feeding 4 level • Narrow Band Imaging (NBI) for capillary pattern visualization• Compatibility with existing equipment Video system center:• Color CCD system • White balance adjustment• HDTV compatibility – 1.080 effective scanning lines• Narrow Band Imaging (NBI) for capillary pattern visualization• Image size selection, 3 steps • Image enhancement settings, structural and edge enhancement, 4 levels • Automatic gain control• Electronic magnification, 2 steps• Memory card storage (including memory card 2 psc. – 1 GB)• “PiP” picture in picture function • Digital outputs (IEEE 1394, Fire-wire, DV,....) • Iris control, 2 modes (peak and auto) • Freeze control of image • Patient data, scope identification • Scopes switches (4 pcs.) for functions • Adjustment of colors: Red, Blue,...(±8 steps) Trolley – Mobile workstation:• Sliding keyboard tray• LCD Monitor arm• Scope hanger• Universal storage container Color Monitor:• LCD Monitor• 26 inch• HDTV Resolution 1920x1200 (WUXGA)• Medical grade• Compatible 5:4 and 16:9• Picture in Picture function, Picture on Picture function, Rotate, Mirror• Contrast ratio 1000:1• View angle 178°• Inputs: HD/SD/SDI (2x), Video, Y/C, Analog 2 x RGB (4 x BNC), DVI (2x)... Suction pump:• Nominal vacuum: 0.85 bar• Nominal free air flow rate. 20 l/min. • Variable suction• Suction bottle 2 l• Set of filters for suction pump (10 pcs.) • Set of spare tubes for suction pump (≈50 pcs.)
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details