## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 141 | Bidder: |  | TOTAL DAP PRICE OF LOT No141: |  |  | 0
Delivery address: Institut za nuklearne nauke `Vinča, Mihajla Petrovića Alasa 12-14, 11351 Beograd-Vinča, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Instrument for analysis of fuels |  |  | 1 | 0 | 0
| 4279 | Termogravimetric analyzer Analyzer for determination of weight losses caused by temperature increase, as follows: moisture, volatile compounds, etc., and ash determination in organic materials using a controlled resistance furnace with selectable inert or oxidation atmosphere. Complying with AOAC, AACC, and ASTM methodology. Technical specifications of the termogravimetric analyzer: Sample size: Macro sample sizes (50 mg to 5 g) Number of samples: Up to nineteen samples (+ 1 reference) during one analysis and with possibility to have up to 38 samples using a dual-furnace configuration Carousel: Ceramic, not susceptible to warping at temperatures up to 1000°C Carousel mechanism: Pneumatic for accurate crucible placement Crucibles: Ceramic, capable of holding macro sample sizes (50 mg to 5 g)Balance: Integrated four-place, shielded and protected from any interference from other internal components Balance resolution: 0.0001 g Precision: ≤0.02% standard deviation at various temperatures, flow rates, and atmospheres (1 σ, 1 g inert sample) Temperature range: From ambient to 1000°C Temperature control: Accuracy: 2% of setpoint, or ± 2°C; whichever is greater Stability: 2% of setpoint, or ± 2°C; whichever is greater Temperature ramp rate: Variable: Ambient to 104°C: up to 15°C/minute 104°C to 1000°C: up to 50°C/minute Thermometric system: Two pieces of thermometers built in furnace with mathematic model for calculation of temperature inside of crucible Sample preparation system: Mill for sample preparation Safety protection: Dedicated over-temp protection thermocouple PC System: Operating software, monitor, color laser jet printer Operating software: Windows®-based, with compliance to 21 CFR Part 11 regulations, containing an on-board help manual and with control up to two separate furnace units Service diagnostics: Remote service diagnostics with a separate software package that is designed specifically to connect the user with the Instrument vendor’s service personnel Service accessibility: From the front of the instrument Environmental conditions: Operating temperature: 15°C to 35°C Humidity: 20% to 80% non-condensing
2 | Instrument for ultimate analysis of fuels |  |  | 1 | 0 | 0
| 8773 | Elemental Analyzer CHNS Analyzer for simultaneous determination of carbon, hydrogen, nitrogen and sulphur (with independent add on module) in organic and most of inorganic matrices (solids and liquids). Instrument utilizes a combustion technique independent detectors optimized for each element (thermal conductivity detector for nitrogen and three infrared detectors for: carbon, hydrogen and sulphur). Measurement of all elements of interest, speeding the analysis time and eliminating the need for combustion gas separation techniques. Complying with AOAC, ASTM, ISO, AACC, AOAC and ASBC methods of analysis.Technical specification of Elemental Analyzer CHNS:Instrument range: Carbon: 0.02 to 175 mg carbon Hydrogen: 0.1 to 12 mg hydrogen Nitrogen: 0.04 to 50 mg nitrogen Sulphur: 0.01 to 20 mg sulphur Precision range: Carbon: 0.01 mg or 0.5% RSD, whichever is greater Hydrogen: 0.05 mg or 1.0% RSD, whichever is greater Nitrogen: 0.02 mg or 0.5% RSD, whichever is greater Sulphur: 0.005 mg or 1.0% RSD, whichever is greater Sample mass: CHN: up to 250 mg S: up to 350 mg Readibility: 0.0001 % Analysis time: 4.5 minutes Furnace CHN: Two-stage (primary and anfterburner) combustion furnace, up to 1050°C, with independent temperature control, accuracy of ±1% of the set point Providing 100% oxygen environment with additional oxygen supplied above the sample using a quartz lance Furance S: Horizontal resistance type, up to 1450°C, accuracy of ±1% of the set point Providing 100% oxygen environment with additional oxygen supplied above the sample using a quartz lance Moisture removal: With pre-chiller and thermoelectric cooler Detection: 4 independent detectors for measurement of combustion gases (CO2, H2O, SO2 and N2):- 3 non-dispersive infrared absorption detectors (IRD): for C, H and S - one thermal conductivity detector (TCD) for N System without any form of chromatographic separation or whole gas analysis.Calibration: Single-point; Multi-point (linear, quadratic, cubic polynomial)Autoloader: 30-position, low-maintenance, rotating carousel, optionally upgradeable to 120 positions PC System: Operating software, monitor, color laser jet printer Operating software: Windows®-based, with compliance to 21 CFR Part 11 regulations, containing an on-board help manual, with wellness checks of the instrument based on user-defined service conditions, real-time service diagnostics, including ambient charts of instrument temperatures, pressures, and detector signal; manual control of solenoids and switches; automated leak checks; and network and communications diagnostics Automatic system check: Verifies internal network communications, solenoids and switches, system pressures, furnace control, and temperature Service diagnostics: Remote service diagnostics with a separate software package that is designed specifically to connect the user with the Instrument vendor’s service personnel Environmental conditions: Operating temperature: 15°C to 35°C Humidity: 20% to 80% non-condensing
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F9: Insert CIF price
Note on D10: FILL all white fields, write brand, trade name, model and all relevant trade details