## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 170 | Bidder: |  | TOTAL DAP PRICE OF LOT No170: |  |  | 0
0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Thermal Properties Analyzer |  |  | 1 | 0 | 0
| 345 | Thermal conductivity measuring instrument for laboratory use Thermal Conductivity Meter – instrument for analyzing the thermophysical properties of various materials (thermal conductivity and heat transmission resistance of building, insulation and other materials). Easy to use, already calibrated or calibration material inclueded. Most important requirements: Measuring range of thermal conductivity (minimum): 0.02 – 1.0 W/m•K Temperature Range (minimum): 0º to 50ºC Accuracy: Better than 5% Power requirements: 220 V Portable apparatus for in-situ measurement of thermophysical properties Portable measurement apparatus for analysis of thermal resistance and thermal transmittance of building elements, by in-situ measurement. Measurements according to ISO 9869 and ASTM C1155 / C1046 standards. Heat flux sensor and thermocouples are important to be equiped with longer cable (atleast 20m). Primary use for analysis of buiding envelopes and walls. The apparatus should be already calibrated or calibration material should be inclueded, as the adequate software for readout. Delivery address: Građevinski fakultet u Subotici, Kozaračka 2/a, 24000 Subotica, Serbia
2 | Xenon Flash Thermal Constant Analyzer |  |  | 1 | 0 | 0
| 1473 | Thermal Constants Analyzer for contact free nondestructive testing of thermal diffusivity and thermal conductivity of solids and liquids in wide temperature range. Technical specifications 1. Temperature range: from room temperature to minimum 500oC 2. Pulse source, pulse energy and sensor type: xenon flash lamp for generation of irradiating energy pulse up to 15J/pulse with mercury cadmium telluride (MCT) detector or InSb detector for detection of temperature rise 3. Furnace: Resistance furnace with automatic furnace lift 4. Heating and cooling rate: 0.01 up to 50°C/min 5. Vacuum: up to 10E-5mbar 6. Atmosphere: inert, oxidizing or reducing 7. Thermal diffusivity measuring range: 0,01 up to 1000 mm2/s 8. Thermal conductivity measuring range: 0.1 up to 2000 W/mK 9. Samples and sample holders: Various sample holders for different sample geometries from 10 mm square to 30 mm round samples, up to 6 mm thickness, with maximal sample number depending on sample size (up to 6). 10. Accuracy:  3% 11. Electronics for data acquisition and complete data analysis software package 12. USB or Ethernet interface 13. Included installation and training 14. Included packing and L/C charge Delivery address: Tehnički fakultet u Boru, Vojske Jugoslavije 12, 19210 Bor, Serbia
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F9: Insert CIF price
Note on D10: FILL all white fields, write brand, trade name, model and all relevant trade details