## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 174 | Bidder: |  | TOTAL DAP PRICE OF LOT No174: |  |  | 0
Delivery address: Institut za nuklearne nauke `Vinča`, Mike Petrovica Alasa 12-14, Belgrade, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Termomechanical analyzer Q400EM |  |  | 1 | 0 | 0
| 5021 | Instrument Performance:- Temperature range: -150 to 1000*C (with temperature precision * 1*C) - Heating rate range: 0.1 to 100*C/min - Isothermal temperature control of *0.1*C - Furnace cool-down time: < 10 minutes from 600*C to 50*C - Measurement Precision: * 0.1% - Sensitivity: 15nm - Resolution (digital): 0.15 nm - Dynamic baseline drift: <1 μm (-100 to 500*C) - Force range: 0.001 to 2N (with force resolution 0.001 N) - Frequency range: 0.01 to 2 Hz - Atmosphere: Inert, oxidizing, or reactive - Maximum sample size(solid): 26 mm(L); 10 mm(D) - Maximum sample size (film/fiber): 26 mm(L); 0.5 mm (T); 4.5 mm(W) TMA must have the following features: - Automated furnace movement, probe movement, probe zeroing, sample length measurement, force loading. Plus local control at the module including start/stop, furnace position, probe movement, zeroing, sample length and real-time display of sample temperature and experiment status. - Ability to automatically open furnace at the end of run, raise probe off the sample, plus ability to automatically terminate a run upon reaching a predetermined value. - Ability to use compressed air/nitrogen cooling. Cooling from 1000*C to 50*C in 10 minutes or less. - Ability to enter up to five points for temperature calibration. - Ability to store calibration constants in module memory. - Ability to generate high quality data in the compression, tension, and 3-point bending modes of deformation. - Ability to include sample geometry factors to compute rheological signals. - Ability to generate temperature steps/ramps at constant strain and constant stress, and stress and strain steps/ramps at constant temperature, plus force steps and ramps at constant temperature. - Ability to generate tangent or secant moduls plots versus stress, strain, time or temperatrure.- Ability to perform creep and stress relaxation measurements, with capability to compute data in rheological units (% strain, creep compliance, recoverable compliance, relaxation modulus). - Ability to perform dynamic TMA (DTMA) tests. - Ability to perform modulated TMA (MTMA) tests. TMA must offer the following probe and fixture configurations: - Standard and macro expansion probe - Penetration probe - Tension accessory kit for films and fibers - Dilatometer accessory for cubical coefficients of expansion - Parallel plate accessory kit - 3-point bending kit
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details