## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 178 | Bidder: |  | TOTAL DAP PRICE OF LOT No178: |  |  | 0
Delivery address: Tehnološki fakultet u Leskovacu, Bulevar oslobodjenja 124, 16000 Leskovac |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Inductively Coupled Plasma Optical Emission Simultaneous Spectrometer (ICP-OES system) |  |  | 1 | 0 | 0
| 8299 | Optical range 130-770 nm or better Analytical view Axial view Optical system Paschen-Rounge or Czerny-Turner or Echelle Plasma generator „free running” 27 MHz Power generator stability ≤ 0.2 % Peristaltic pump 4 channels, fully computer controlled Optical Resolution CONSTANT resolution over the whole range: 0.010 nm or better in UV and VUV range, and 0.015 or better in VIS range Optic stabilisation Over 100C Gas flow Fully computer controlled Salt tolerance ≥ 20% salt concentration Total Argon consumption during measurement 14 l or better Measurements Simultaneous measurements of major, micro and traced elements Recording Recording and memorising of whole spectra not only results, post analysis reprossesing–recalculation without new measurements Recalibration Fast recalibration of the system using only 1 reference Power generator Adjustable power generator from 0.7 to 1.7 kW or better Limits of detection ≤ ppb level Possibility of measuring NONMETALS B, P, S, Cl, I and Br Dinamic linear range 6 range of magnitude or better Autosampler ≥ 50 places Software Easy to operate, intuitive Data system specifications PC with monitor TFT 19" and laser printer. Full control of all ICP-OES and Autosampler parameters
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details