## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 8 | Bidder: |  | TOTAL DAP PRICE OF LOT No8: |  |  | 0
0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Gas chromatrography |  |  | 1 | 0 | 0
| 643 | Gas Chromatograph Specification with FID and ECD and autosampler – 1 unit 1. Gas Chromatograph specification: Temperature range: From 4 °C above ambient up to 450 °C Oven ramp rates: 2 °C/sec up to 110 °C Oven ramps/plateaus At least 15 ramps and 16 plateaus Split/splitless inlet – 2 pcs Split/Splitless capillary GC inlet with bac-pressure control, pressure range up to 100 psi, pressure presicion 0.001 psi Max temperature – up to 400°C Flame ionization detector – 1 pcs Flame ionization detector with electronic pneumatic control for all detector gases – Air, H2, and make-up gas. Linearity – up to 107 Minimum limit of detection < 1.5 pg C/s and better Data rate: up to 500 Hz Electron capture detector – 1 pcs ECD Linear dynamic range: > 5 × 104 with lindane ECD Minimum detectable level:< 6 fg/mL lindane Diagnostic and GLP feature Extensive diagnostics, error detection and reporting through GC software; continuous tracking of instrument usage with pre-defined and user settable limits and feedback messages. Electronic records of maintenance and errors. 2. Autosampler specification: Vial capacity Up to 294 x 2ml viala; up to 96 10ml/20ml viala Injection mode Liquid and headspace Syringe volume Liquid from 1.2 µl - 500 µl Headspace from 1ml – 5 ml Software control Fully integrated into GC software 3. Software for full instrument control including autosampler, data analysis and report generation. 4. Brand name PC with monitor and laserjet printer according to the manufacturers specification and recommendation. Delivery address:Univerzitet Edukons, VOJVODE PUTNIKA bb, 21 208 SREMSKA KAMENICA
2 | Portable Gas Chromatograph |  |  | 1 | 0 | 0
| 1669 | Fully transportable GC/SAW designed to speciate and quantify organic compounds. Supplied with rechargeable lithium ion battery. Analysis Time 5 – 60 seconds Recycle Time 30 seconds minimum Precision and Accuracy Standard Deviation <2% Sensitivity Can analyze vapors in 10 seconds with sensitivity in parts-per-billion for many compounds. Detects hydrocarbons in the range of C4 – C25. Sensitivity will vary by compound sampling time, matrix, interferences and detector temperature settings. Sampling Internal Sample Pump Sample Introduction: ~.5 ml/second Time programmable : 1-300 seconds Compound Identification Automatic with user calibration Communications RS-232 between controller and 4300. Bluetooth or RS-232, user selectable Dimensions Head Weight 5.7 lbs 2.6 kg Length 12.5 in 31.8 cm Width 4.3 in 10.9 cm Height 6.8 in 17.3 cm Support Weight 18.7 lbs 8.5 kg Length 12.5 in 31.8 cm Width 9.7 in 26.4 cm Height 5.8 in 14.5 cm Charger Weight 7.7 lbs 3.5 kg Length 13.5 in 34.25 cm Width 9.7 in 14 cm Height 3.7 in 9.5 cm Environmental Conditions Operating Temperature Range: 32oF to 105oF (0oC to 40oC) Relative Humidity: 0 – 95% noncondensing Power Battery Pack: 28VDC, 16 AHr – Lithium-Ion (5 hr use typical) Charger Power 100 – 127 VAC at 3 amps – 50/60Hz 200 – 240 VAC at 1.5 amps – 50/60 Hz Column Limits 35oC to 225oC – depending on column Column Ramping Ramped from 0 – 18oC/second Carrier Gas Helium (Min. 99.999% Purity, #6) Typical usage is 300 tests per day on one helium charge Helium (Replaceable Cylinder), 95cc at 17.6 MPz (2560 psi).Typical usage: 200-300 tests Detector Surface Acoustic Wave (SAW) Quartz microbalance Dynamic Range: 2 x 105 Temperature: 0oC to 150oC, programmable Inlet Connection: Stainless Steel LUER inlet port Temperature: 50oC to 200oC System Controller Software Intel Pentium 100 MHz or better processor Minimum: 16MB RAM 1GB Hard Drive Microsoft Windows (any version) Hardware - Vapor Analyzer Hardware - Tools for Vapor Analyzers Sampler Injector Kit (includes Alkane Solution) Accessories System Controller (with RS232 Serial Port). Preloaded with MicroSense User kits User Tool Kits Supplies Columns (VOC) - DB-624 and DB5 Saw Detector replaceable cylinders (100 cylinders) Cases Custom Rolling Case Training and services Training at Company for several days including travel and accomodation of our researcher Delivery address:Institut za biološka istraživanja `Siniša Stanković` u Beogradu, Bulevar despota Stefana br. 142, 11 000 Beograd
3 | Gas chromatograph |  |  | 1 | 0 | 0
| 4276 | GC Oven Max Oven Temperature 450o C Temp Ramps 32 ramps with 33 isothermal holds Cool - Down Rate 450oC to 50oC in 4 min Maximum achivable temperature ramp 125 °C/min Injector S/SL Split Ratio Up to 12500:1 Maximum Temperature 400° C Column suitability Capillary columns (50 μm to 530 μm i.d.)Inj / GC detectors max Upgrade Options Possibility (in sum) up to 2 Injectors and 3 Detectors (one of them mass spectrometer described below). Modular injector and detectors and fast changing of injector and detecor modules. Carrier Gas Control Carrier gas control precision 0.001 psi Mode Constant and programmed pressures and flows Pressure Range 0–1000 kPa Mass spectrometer Analyzer Single Quadrupole EI Scan Sensitivity 1ul of 1pg/ul OFN produces S/N 450:1 for m/z 272 in scanning range 50-300u with helium as carrier gas Positive CI Scan Sensitivity 1µL of 100pg/µL benzophenone will give 300:1 RMS S/N at m/z 183 Negative CI Scan Sensitivity 2µL of 100 fg/µL ctafluoronaphthalene (OFN) will give 600:1 S/N at m/z 272 Source EI and CI in one source. Changing beetween ionization techniques (EI/CI) without removing source. Easily maintenance without breaking a vacuum, proceudure should take time in several minutes Ionization EI with full scan, SIM, and sequential full scan/SIM. Chemical ionization, including PCI, NCI scans in a single injection. EI/CI combination.Electron Energy Range 0-150 eV Emission current up to 350uA Filament 2 filaments Mass filter Dual stage mass filter with off-axis ion guide Mass Analyzer Mass range 1.5-1050 u with unit mass resolution or better Scan Features Scan rate (amu/sec) Better than 11.000 u/sec Full Scan YES SIM YES Segment Scan Up to 10 Detector General discrete dynode electron multiplier and electrometer Linearity 0-60µA or better Pump Pumps Turbomolecular pump more than 230l/s He and machanical pump Reagent Gas Software-switchable dual reagent gas with digital flow control Autosampler Sample capacity 105 samples Vial capacity 2ml / 105 vials Injections/vial 0-99 Injection volume 5 μL Syringes Standard: 10 μL Syringe Rinsing Pre and/or Post injection, Solvent selection (Single or combined mode), Sample pre-washes, Bubble elimination SOWTWARE and PC with monitor Complete computer control of parameters for GC, mass spectrometer and liquid autosampler. Real time data analysis. Software, specialized for environmental laboratories. Latest realesed NIST library. Complete audit trail for GLP and QC purposes-records on who changed what, when and why. Brand name PC with flat monitor. Delivery address:Institut za nuklearne nauke `Vinča`, Mihajla Petrovića Alasa 12-14, 11351 Beograd-Vinča
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F9: Insert CIF price
Note on D10: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F11: Insert CIF price
Note on D12: FILL all white fields, write brand, trade name, model and all relevant trade details