## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 11 | Bidder: |  | TOTAL DAP PRICE OF LOT No11: |  |  | 0
0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Gas chromatography with mass detector and auto sampler for liquid samples: GC/MS |  |  | 1 | 0 | 0
| 3675 | Flow mode: constant pressure and linear velocity Oven: Temperature range: room +5 to 450˚C Temperature program stage: up to 20 temperature ramps Temperature speed: better than 120˚C/min Cooling rate: less than 4 min for cooling from 450˚C to 50˚C Injector: Split/splitless injector with advance flow control Temperature range: room +5 to 450˚C Mass selective detector: Ionization mode: EI (electron impact) Saving gas and electrical consumption Option: capable for NCI/PCI upgrade Mass filter: Metal quadruple with pre-rods/filters Mass range: from 1, 5 up to 1090m/z Mass Cycle Scanning Speed (single scan): > 10.000 u/sec Sensitivity Scan (EI) (OFN): >S/N 400 ：1pg Filament: Double (EI, PCI, NCI) (Automatic switching in analyzing） Main Pump: Turbo molecular Ion source temperature: up to 320˚C Transfer line temperature: up to 350˚C Software Should be Windows compatible. Instrument control and data collection Simultaneously data collection and collected data processing Auto sampler for liquid samples Injection volume 0,1uL to 250uL Capacity: 15 vials, 1,5 or 2ml volume Min 100 pcs. 1,5ml vials Commercially available library with mass spectra (Wiley, NIST….) PC, printer with OS with license. Installation and customer education should be included. Delivery address:Institut za veterinarstvo u Novom Sadu, Rumenacki put 20, 21000 Novi Sad
2 | Gass Chromatograph / Mass Spectrometer |  |  | 1 | 0 | 0
| 3731 | Gas Chromatograph specification:Temperature range: From 4 °C above ambient up to 450 °C Oven ramp rates: 2 °C/sec up to 110 °C Oven ramps/plateaus At least 15 ramps and 16 plateaus Combined inlet – 1 pcs Should have possibility to work as split/splitless inlet and to be able to inject large sample volume .Mode of operation -cold/hot split/splitless -pulsed split/splitless -solvent vent -direct injection Temperature range: from 10ºC above room temperature up to 450 ºC.Number of programmed ramp: at least 9 ramp with heating rate of 900 ºC/min. Pressure operating range: Up to 100psi, with 0.001 psi pressure step Diagnostic and GLP feature Extensive diagnostics, error detection and reporting through GC software; continuous tracking of instrument usagewith pre-defined and user settable limits and feedback messages. Electronic records of maintenance and errors. Autosampler specification:Vial capacity Up to 150 x 2ml vials Injection volume 0.2 µl - 500 µl Software control Fully integrated into GC software Mass Spectrometer specification:Mode (standard): EI Ion source: Noncoated inert EI source; Ion source with two filament design Ionization energy: user selectable 5-240 eV Ionization current: 0-315 uAmps Number of independent thermal zones and ranges: 3; transfer line temperature temperature: ambient to 300 °C; source temperature: up to 300 °C; quadrupole temperature, up to 200°C Mass filter: Monolithic hyperbolic quadrupole Mass filter protection: Mass range: Entrance lence 2 – 1000 amu Detector: EM (Electron multiplier) with replaceable horn Dynamic range (electronic): 10e6 Scan rate (electronic): 11000 amu/s with 0.1 amu step size and better SIM: 60 ions x 100 groups Vacuum pump: turbomolecular pump, accepts up to 2 ml/min helium Maintenance access: Source, filaments, lenses, mass filter, and detector on removable plate Instrument control: Data system and local user interfacethat allows for operation of system (tuning, venting, start/stop, loading of methods) without interfacing with data station Full scan sensitivity, EI mode: 1pg octafluoronaphthalene shall yield a signal-to-noise ratio of at least 600:1 RMS at nominal m/z 272 ion with Helium as carrier gas, with hydrogen signal-to-noise ratio of at least 100:1 Linear dynamic range over 4 orders of magnitude for calibration curves (quantitation) Software for full instrument control including autosampler, data analysis and report generation. MS Library NIST and Wiley latest edition. Brand name PC with monitor and laserjet printer according to the manufacturers specification and recommendation. Delivery address:Tehnološki fakultet u Novom Sadu, Bulevar Cara Lazara 1, 21000 Novi Sad
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F9: Insert CIF price
Note on D10: FILL all white fields, write brand, trade name, model and all relevant trade details