## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 59 | Bidder: |  | TOTAL DAP PRICE OF LOT No59: |  |  | 0
Delivery address: Institut za prehrambene tehnologije u Novom Sadu, Bulevar cara Lazara 1, 21000 Novi Sad, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | XDS analyzer |  |  | 1 | 0 | 0
| 9050 | The device for the development and application of near infrared spectroscopy within the spectral range 400-2500 nm to examine the various components in the samples in the solid and/or liquid state (granular, powder, paste, liquid). TECHNICAL FEATURES: - Dispersive monochromator - Reflectance measurement - Rapid scanning distortion correction - Rapid offset grating for focus improvement - Wavelength range: 400 – 2500 nm - Dual detector system: Silicon (400-1100), Lead Sulphide (1100-2500) - Data acquisition rate: 2 scan/sec - Spectral resolution: 2 nm - Wavelength accuracy: <0,05nm - Photometric, wavelength and bandpass response for the instrument must be established by the NIST traceable standards and methods - Operating temperature: 5-35°C - Operating humidity: 10-90% relative - Factory standardized hardware, that enables seamless calibration transferability between instruments of the same type - Instrument operation software fully compatible with WinISI software - Horizontal position of the sample, light source and detector bellow the analyzed sample - Mini sample cups for homogeneous samples - Coarse granular sample cell for reflectance analyses of inhomogeneous samples - Liquid sample kit - NIR reflectance standard for instrument standardization and performance monitoring The producer is obliged to provide: - Installation - Operator training at customer site, immediately after the installation and confirmation of equipment specification - Instruction manual - Application, technical and service support in Serbia
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details