## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 67 | Bidder: |  | TOTAL DAP PRICE OF LOT No67: |  |  | 0
Delivery address: Fakultet tehničkih nauka u Novom Sadu, Trg Dositeja Obradovića 6, Novi Sad, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | LIDAR - SPECTRAL CAMERA HS V8E V10E (A) with FieldSpec Pro FR Portable Spectroradiometer (B) |  |  | 1 | 0 | 0
| 446 | The item no. 446 consists of two parts: 1. Hyperspectral camera and 2. Portable Spectroradiometer for camera calibration In the remaining text, specifications for two parts are given separately, as they can come for two different manufacturers.1. Hyperspecteral camera:The following characteristics are REQUIRED:Optical characteristics• Spectral range: 400-1000 nm • Spectral resolution (30 μm): 22.8nm • Spectral sampling: 0.72 - 5.8 nm/pixel • Spatial resolution: RMS spot size <9μm • Aberrations: Insigniﬁcant asligmatism, keystone or smile • Numerical aperture: F/2.4 • Slit width options: 30 μm (18, 50, 80, 150 μm) • Eﬀective slit length: 11.84 mm • Total eﬃciency (typical): > 50% independent on polarization • Stray light: < 0,5% /halogen lamp, 590nm LPF) Electrical characteristics • Sensor: interline CCD • Pixels in full frame: 1600 (spatial) x 1200 (spectral) • Active pixels: 1600 (spatial) x 840 (spectral)• Pixel pitch: 7.4 μm • Camera output: Digital 12 bit • Interface: Base camera link • Camera control: Camera link • Frame rate: 33 fps (full frame) up to 120 fps (ROI) • Exposure time range: 0.1 - 100 ms • Power consumption: < 5W • Input voltage: 24V Mechanical characteristics • Cased • Body: Anodized aluminium with mounting screw holes • Lens mount: Standard C-mount • Shutter : Yes, with USB control 2. Portable Spectroradiometer for camera calibration The following characteristics are REQUIRED:• Spectral range: 350 - 2500 nm • Sensors: One 512 element photo diode array and two thermoelectrically cooled, "graded index", extended range InGaAs photodiodes • Sensor linearity: +/-1% • Dispersion elements: One fixed and two fast scanning holographic reflective gratings.• Sampling interval: 1.4 nm for 350 - 1000 nm, 2 nm for 1000 - 2500 nm.• Spectral resolution: 3 nm @ 700 nm, 10 nm @ 1500 nm, and 10 nm @ 2100 nm.• Wavelength accuracy: +/-1 nm.• Scan time: A new spectrum generated every 0.1 seconds for the entire spectral range.• Data storage: Data are stored in binary form on the built-in hard disk drive (included in the notebook with battery that comes with the device).• Fiber optic cable input • Spectrometer battery pack included • Software for data acquisition and storage of Raw DN, reflectance, radiance and irradiance measurements included. • Possibility for real-time calculations of reflectance.• Real-time display of Raw DN, reflectance spectra, radiance or irradiance.• Real-time spectrum averaging for up to 31,800 spectra for increased signal-to-noise.
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details