## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 72 | Bidder: |  | TOTAL DAP PRICE OF LOT No72: |  |  | 0
Delivery address: Institut za nuklearne nauke `Vinča`, P.O. Box 522, 11001 Belgrade, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Thermoluminescent Dosimetry reader |  |  | 1 | 0 | 0
| 1543 | A TLD reader-analyzer (a linear, programmable heating system and a cooled photomultiplier tube with associated electronics to measure the TL light output) for loose TLD materials (ribbons (chips), rods, micro-cubes) and TLD powders; Planchet heating incorporates welded thermocouple or highly stable platinum alloy heater; Three-stage linear programmable heating profile includes pre-heat (initial heating), acquire (reading stage) and anneal (terminal heating) cycles. Duration of each stage up to 60s; Heating temperature capability up to 600 °C with accuracy 0.1°C; 7 decade glow curve acquisition range or dynamic range of the I-F converter: 8 decades; Supplied with operational software; Optional neutral density filters to extend the high measurement range; Built-in optical filter for elimination of infra-red radiation; Argon or nitrogen shield recommended for low intensity TL doses and elimination of chemical luminescence; Thermoelectric PMT cooler for maximum gain stability; Stabilization of HV PMT power supply with accuracy of 1/10000; Automatic background subtraction capability; The results of measurements and the parameters of settings of the equipment should be able to be transmitted to a PC system and printer by an adequate interfaces and exported in ASCII format
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details