## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 121 | Bidder: |  | TOTAL DAP PRICE OF LOT No121: |  |  | 0
Delivery address: Institut za medicinska istraživanja u Beogradu, Dr Subotića 4, 11129 Belgrade, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | nitric oxide analyser |  |  | 1 | 0 | 0
| 9401 | The instrumentation utilises ozone chemiluminescence technology. The analyser includes a vacuum pump for its operation. The instrument should be capable of measuring nitric oxide in both liquid and gaseous samples. This includes measuring nitric oxide in human exhaled breath and nasal samples, both on-line and off-line. Sensitivity: at least 0.5ppb NO in gas phase and 1 picomole in a liquid sample. Measurement range: up to 500ppm NO in gas phase and up to 1 milliMole NO in liquid phase. Response time of the instrument: max 1sec for 90% signal rise, max 1sec lag time. Outputs: digital RS-232 output (at least 38400 baud) and an analog output (0-1V or 0-10V). Repeatability of measurement: +/- 5% or better for both gaseous and liquid samples. Gas sampling system for gaseous applications. Glass purge system for determination of S-nitrosothiols, nitrite, nitrate and nitric oxide in liquid samples. Pressure monitor and display for the ozone chemiluminescence reaction cell. Supported software suitable for data acquisiton and data processing for liquid and gaseous applications. To provide service and support for the lifetime of the instrument.
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details