## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 134 | Bidder: |  | TOTAL DAP PRICE OF LOT No134: |  |  | 0
0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Cell Sorter |  |  | 1 | 0 | 0
| 1525 | Flow cytometry cell sorter based on piezzoelectro driven catcher tube mechanism for cell sorting in aerosolfree completely enclosed environment with a cell concentrator module. Sorting purity: more than 95% Capture rate: 300 cells/second Sort modes: single cell (aerosol free), exclusion (aerosol free), recovery (aerosol free) Recovery: more than 50 % Sterile sorting: aerosol free Cell concentration: yes with cell concentrator module Expected price: 38000 euros Delivery address: Hemijski fakultet u Beogradu, Studentski trg 12-14, Belgrade, Serbia
2 | Equipment for flow cytometric analysis – flow cytometer with computer and laser printer |  |  | 1 | 0 | 0
| 2392 | Benchtop research flow cytometer, minimum 2 Lasers, 8 Parameters (6 colors plus FSC/SSC) Optics • Minimum two-lasers (488 nm and 640 nm) supporting 8 parameters – 6 colors plus FSC/SSC • Aautomated laser alignment Fluidics • Not pressurized, open system that allows fluids sddition during measurement, and thereby measuring kinetics (e.g. Calcium-Flux) • Wide range of input devices in the manual loading position, including 12 x 75-mm tubes made of polystyrene, polypropylene, or K-resin, 15-mL and 50-mL conical tubes, and 1,500-µL or 500-µL microcentrifuge tubes • High sensitivity mode - for better resolving dimly expressed populations • Flow sensor - No requirements for beads for absolute counts, cell count, no invasive, no moving parts, lifetime calibration Computer Hardware and Software • Software with pre-defined research assays for apoptosis, cell cycle, cell proliferation, and cytokine detection • allows multitasking and parallel working (data acquisition and analysis in parallel) • Universal Settings - No need for daily compensation Delivery address: Insitut za virusologiju,vakcine i serume `Torlak` u Beogradu, 458 Vojvode Stepe, Belgrade, Serbia
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F9: Insert CIF price
Note on D10: FILL all white fields, write brand, trade name, model and all relevant trade details