## Sheet1
LOT 1 | Bidder: |  | TOTAL DAP PRICE OF LOT No1: |  |  | 0
0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | lvaX - X ray (EDXRF) fluorescence spectrometer |  |  | 1 | 0 | 0
| 434 | Svrha: uređaj omogućuje detekciju i određivanje koncentracije svih hemijskih elemenata u periodnom sistemu, u rasponu koji počinje (računajući od lakšeg ka težem elementu) minimalno od Cl – hlora (atomski broj 17), a po mogućnosti od još lakšeg N - azota (atomski broj 7) pa sve do U- uranijuma (atomski broj 90). Vrsta uzoraka: - biljni uzorci (različiti delovi biljaka kao npr. uzorci listova, stabla, kore, korenova, plodova, semena), uzorci zemljišta, uzorci vode (tečnosti) Granica detekcije: - poželjno je da se elementi mogu čitati minimalno u milionitim delovima mase (<ppm – parts per milion), a poželjno da je rezolucija jača i da se čita i u milijarditim delovima mase (ppb – parts per bilion), po jediničnoj masi uzorka. Opis aparature: -glavna konzola za analizu sa X-ray izvorom pobuđivanja, i odgovarajućim detektorom i selekcionim sistemom koji omogućava analizu navedenih uzoraka i zadate detekcione granice - odgovarajući pribor za obradu i pripremu uzoraka koji ide u paketu sa instrumentom - odgovarajući „autosampler“ (automatizovani sistem unošenja i analiziranja uzoraka) koji je moguće dobiti uz izabrani aparat, po mogućstvu za što veći broj uzoraka biljnog materijala, tečnosti i zemljišta - prateća računarska podrška, kod pojedinih modela podrazumeva i računar kao kontrolnu jedinicu sa odgovarajućim softverom, ili kod nekih modela, softver i procesorska jedinica mogu biti integrisani u aparat. - odgovarajući „standardni uzorci“ za kalibraciju koji idu u paketu sa instrumentom Delivery address:Prirodnomatematički fakultet u Novom Sadu, Trg Dositeja Obradovića 2, 21000 Novi Sad
2 | XRF spectrometer |  |  | 1 | 0 | 0
| 622 | Analytical method: X-ray spectrometer Method of measurement: Energy dispersive X-ray analysis Acceptable sample type: Solid Liquid Powder Element range: Sodium (Na11) to Uranium (U92) Applications areas: Environmental Biological Food Soil Sediment Geological samples Concentration range: ppb to 100% Sample chamber: Atmosphere: Air, Vacuum, Helium 8 to 15 position automatic sample changer Detector: High performance Silicon Drift Detector (SDD) or Si (Li) Solid State Detector (SSD) Optimum balance of spectral resolution and high count rate Large active detection area User Interface: External PC computer system Operating system Keyboard and mouse LCD Monitor Printer Shipment: Delivered to Belgrade Installation: Installation by authorized personnel Training tuition: Include Warranty: 12 months from date of shipment Delivery address:Tehnološko-metalurški fakultet u Beogradu , Karnegijeva 4, Beograd
3 | Handheld XRF |  |  | 1 | 0 | 0
| 872 | Tube: 50kV Ag anode X-ray Tube Detector: Geometrically Optimized Large area Drift Detector Intergrated camera Two rechargeable 6-cell battery packs Battery charger AC power supply Carrying case Shielded belt holster SpectraView element scanner Integrated flip touch screen with large intuitive color icons Virtual keyboard for data entry Password-protected set-up and operation with several other radiation safety features 11. Software suite for easy data downloading and viewing, remote control and grade library modification PC connection cable Reference samples and/or certified reference material (s) Bluetooth Wireless Connectivity RFID technology for test stand recognition Additional Prolene aperture windows for light-element Analysis Mobile Test Stand Bulk Sample Kit for sample collection, preparation and analysis Mode: Soil analysis; Metal alloy analysis package including Chemistry, Signature Match, and SuperChem modes; direct analysis of precious metal Small spot Collimation: Colimates x-ray exitation beam to a 3mm spot size Global Sat: Bluetooth GPS Delivery address:Filozofski fakultet u Beogradu , 11000 Belgrade, Serbia
4 | Portable XRF |  |  | 1 | 0 | 0
| 8588 | Standard hardware: 50kV/200uA, 2 Watt Ag Anode X-Ray Tube Geometrically Optimized Large Area Drift DetectorDetector. Integrated CCD Camera. Two rechargeable 6-cell battery packs Battery charger; AC power supply SpectraView element scanner Bluetooth Wireless Connectivity PC connection cable Reference samples and/or certified reference material (s) Fundamental parameters based metal alloy analysis package RFID technology for test stand recognition Integrated GPS receiver GlobalSat Bluetooth GPS Able to:  analyze metal alloys;  carry out mining exploration and mapping (metal and mineral content analysis at concentrations >1% in ore and concentrates);  detect soil contaminants (environmental assessments and metals in soils and sediments analysis with total concentrations < 2%);  test electronics and consumer goods for prohibited substances. Detection of the following elements in an SiO2 matrix, a typical soil matrix (SiO2 with Ca/Fe), and SRM matrix: Ba, Sb, Sn, Cd, Pd, Ag, Mo, Nb, Zr, Sr, Rb, Bi, As, Se, Au, Pb, W, Zn, Cu, Ni, Co, Fe, Mn, Cr, V, Ti, Ca, K, Cl, S, P, Si, Al, Mg. Limits of Detection in ppm range. Delivery address:Građevinski fakultet u Beogradu, Bulevar kralja Aleksandra 73, 11000 Belgrade, Serbia
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F9: Insert CIF price
Note on D10: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F11: Insert CIF price
Note on D12: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F13: Insert CIF price
Note on D14: FILL all white fields, write brand, trade name, model and all relevant trade details

## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD