## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 107 | Bidder: |  | TOTAL DAP PRICE OF LOT No107: |  |  | 0
Delivery address: Medicinski fakultet u Novom Sadu, Hajduk Veljkova 3, Novi Sad, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Isokinetic dynamometer |  |  | 1 | 0 | 0
| 2343 | The isokinetic dynamometer should provide a useful method for performance testing and rehabilitation methods for the ankle, knee, hip, wrist, elbow and shoulder. The modes of operation for exercise and testing should include isokinetic, passive, isometric, isotonic and reactive eccentric. Biofeedback should also be provided during each assessment.The isokinetic dynamometer should measure net muscle torque for the following movements:• Ankle (plantar flexion/dorsiflexion; inversion/eversion)• Knee (flexion/extension)• Leg Press• Hip (flexion/extension; abduction /adduction; internal /external rotation)• Wrist (flexion/extension; pronation/supination)• Elbow (flexion/extension)• Shoulder (flexion/extension; abduction/adduction; internal/external rotation) Technical SYSTEM SPECIFICATIONS Features:Multi-Mode operation: Isokinetic, Isometric, Isotonic, Reactive Eccentric and Passive Concentric speed up to 500 deg/sec Eccentric speed up to 300 deg/sec Concentric torque up to 680 Nm Eccentric torque up to 444 Nm Passive speed as low as 0.25 deg/sec - Passive torque as low as 0.7 Nm - Isotonic torque as low as 0.7 Nm Power Requirements Input Voltage 220 VAC Input Frequency 50Hz Clinical Data Station: Windows Operating System Isokinetic Dynamometer Software Flat Panel Touch-Screen Printer Attachments for Ankle, Knee, Shoulder, Elbow, Wrist, Hip Dynamometer Performance Specifications Torque accuracy of +/- 1% of full scale Position range of 330 degrees Position accuracy of +/- 1 degree of rotation
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details