## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 136 | Bidder: |  | TOTAL DAP PRICE OF LOT No136: |  |  | 0
Delivery address: Elektrotehnički institut `Nikola Tesla` a.d. u Beogradu, Koste Glavinića 8a, Belgrade, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | A device for testing the stator winding insulation DCR 50 50kVdc |  |  | 1 | 0 | 0
| 4241 | The DC Ramp test instrument provides a safer off-line test of stator winding insulation compared to a DC or AC hipot test. It can identify problems in the stator winding insulation before an unexpected in-service failure, unlike the pass/fail of a dc or ac hipot test. An automatically ramped or manually stepped direct voltage (up to 60 kV) is applied to the stator winding and the resulting current is measured and plotted against voltage. Examination of the current versus voltage curve allows assessment of stator insulation condition, as well as diagnosis of damage, defects, and deterioration mechanisms. It allows tests to be performed in accordance with IEEE Std 95-2002. • Ground Presence Detection and Protection • Power Input 85-264 V AC, 50/60 Hz • Automatic Ramp or Manual Mode - output voltage and metering up to 60 kV DC • Ramp Rate Control (0-2.5 kV/min) (continuously adjustable) • Output Current 1,0 mA max • Current Metering range switchable, charge and discharge mode • Automatic Trip Protection (high/low tracking, output limit, fault detection) • Built-in Winding Discharge Circuit • Analog and Digital Outputs for Voltage and Current • Connection to PC to: display results as test progresses, store multiple curves and test parameters for later analysis and display curves side-by-side for test comparison • Examination of the current versus voltage curve allows test stoppage before failure of insulation and diagnosis of damage, defects, and deterioration mechanisms. • Compliant with IEEE Standard 95-2002
2 | A device for high voltage impulsed testing of insulation |  |  | 1 | 0 | 0
| 4243 | Features: The test instrumet perform Resistance, HiPot and Surge tests as well as digitize and store data for future use. The Resistance test verifies the existence of dead shorts within the turn-to-turn coils, shows any imbalances between phases due to turn count differences and locates poor wire connections or contacts. DC HiPot testing detects faults in groundwall/earth insulation, and also provides a complete Polarization Index test. The groundwall/earth insulation system consists of the wire and slot liner insulation, wedges, varnish, and sometimes phase paper. Surge testing detects faults in both inter-turn winding and phase-to-phase insulation systems. Using advanced analog-to-digital conversion hardware, it captures the surge test waveform, remembers it, displays it indefinitely and prints it to the included printer. This surge waveform storage capability can be applied to other motors besides simple induction motors. Test instrumet is going to be used for the winding testing of HV rotating mashines – generators, motors, rotor of hydro-generators by storing the wavform from a surge test on one coil, and comparing it to every other coil’s waveform. Test results from up to 10 motors in the field can be stored, retrieved, printed and uploaded to a desktop program for file management and analysis. Each of these 10 records should has its own memory location and each location should be able to store up to three surge wave patterns plus DC HiPot test voltage and current. Surge Test Output Voltage up to 16 kV Maximum Current up to 800A Short Rise time100-200ns according to IEEE Std 522 - 2004 Maximum Energy up to 3J Discharge Capacitance DC Hi-Pot / Megohm Output Voltage 0 - 16 kV Maximum Current 5 mA Micro-Amp Display up to2,000 μA Max Over-Current Trip 5 mA Max Resistance 100,000 MΩ Input Voltage 120 or 220-240 V AC ±10%, 50Hz 4 test clamps (3 phase + 1 neutral), foot switch, warrning lamps Ports: USB, RS 232, conection to PC
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F9: Insert CIF price
Note on D10: FILL all white fields, write brand, trade name, model and all relevant trade details