## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 5 | Bidder: |  | TOTAL DAP PRICE OF LOT No5: |  |  | 0
0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Encapsulation unit VAR D professional |  |  | 1 | 0 | 0
| 855 | Control unit for regulation of vibration frequency, electrostatic dispersion, heating and air pressure Nozzle heating: 30°C to 70°C Vibration frequency: 40 to 6000 Hz Erlectrode tension: 250 to 2500 V Droplet size range: 0,15 to 2,0 mm Dead volume: approx. 2ml Pump rate by air pressure: 0,5 to 200ml/min Max. allowed air pressure in the system: 1,5bar Parts in contact with medium: autoclavable Material in contact with sample: stainless steel, glass, silicone, PTFE Power consumption: max 150W Connection voltage100-240 VAC Frequency: 50/60 Hz Enviromental conditions: for indoor use at temperature of 5 to 40°C Approvals: CE, UL/CSA Accessories: Single nozzles with opening of 0.08, 0.12, 0.15, 0.20, 0.30, 0.45, 0.75 and 1.00mm, made of stainless steel 316L including nozzle rack, Concentric nozzle set for one step core/shell capsule production - 7 external nozzles with opening of 0.2, 0.3, 0.4, 0.5, 0.6, 0.7 and 0.9mm made of stainless steel Pre-filters for nozzle, diameter 7mm (10 pieces) Pressure bottles 500ml and 1000ml,made of glass, with fittings, tubes and air filter, working pressure up to 1,5bar, autoclavable and grounding set. 1 of each Additional requests: The producer is obliged to provide: Installation Operator training at customer site, immediately after the installation and confirmation of equipment specification Instruction manual Local service support by producer certified service technicians Delivery address:Poljoprivredni fakultet u Beogradu, Nemanjina 6, 11080 Beograd
2 | DEA 231 Epsilon – Dielectric Analyzer |  |  | 1 | 0 | 0
| 1367 | one channel dielectric cure monitoring system consisting of: A) Basic Unit with: 1 dielectric sensor, 1 thermocouple measurement, 2 analog input capability; frequency range from 10 Hz to 10 kHz; minimum sampling times: 0.05s; log conductivity range: from -11Si/cm to -4 Si/cm; log Ion Viscosity Range: from 4 to 11 Ωcm Critical Point Control (CPC) software module, 32 bit, for data acquisition, storage and evaluation under MS WINDOWS Features: - automatic identification of four dielectric critical points - event based control at one of the four dielectric critical points - determination of up to 60 critical points associated with dielectric, temperature, pressure, LVDT, and time of day for documentation - tabulated printout or ASCII file export of the calculated values (prerequisite: standard-PC with min. 100 MHz, with serial interface (RS 232), 64 MByte RAM, WINDOWS 95/98/NT4 only and printer) - dielectric extension cable - thermocouple extension cable - dielectric sensor connector B) Computer System C) Accessories and spare parts 1. Furnace for DEA 2. Sensors, Cables and further Accessories D) Software Extension E) Extension for Photocalorimetry Delivery address: Fakultet za fizicku hemiju, Studentski trg, 12-16, p.p 47, 11158 Beograd, PAK 105305
3 | RESEARCH MIXER-SETTLER EQUIPMENT |  |  | 1 | 0 | 0
| 8378 | Equipment ideally suited for: - Development of Reagents used in Solvent Extraction - Evaluation of Process Flow Sheets - Small Scale Production RESEARCH MIXER-SETTLER EQUIPMENT must by compactly designed, self-contained mixer-settler units, built in blocks, one for each function – extraction, scrubbing, stripping, washing etc. The heavier liquid can be recycled. Each unit must by machined from solid material, thus avoiding cemented or welded joints. Solvent extraction laboratory equipment must have: 6 stages extraction and 6 stages stripping. One mixer with speed control units (12 motors for all ). Two pairs of end support. Total flow (aq+org+recycle): 1,5 m/h Surface loading: max. 10 l/h. Pump (2) with flow capacity 7-8 l/h (Aq). Pump (1) with flow capacity up to 15-18 l/h (Org). Material: PVDF (polyvinylidene fluoride) – material must by resists on HCl (6M). Tubings. Control table. The dimensions (LxWxH) of one unit mus by cca. 260x60x240 mm (incl. motor). The active mixer volume is 0,12 liter and the settler volume is 0,48 liter with a loading surface area of 0,006 m2. Recommended max. total flow (org+aq+recycle) is 10 l/h at normal settling velocity. The units in PVDF therefore possess an excellent chemical resistance to liquids (specialy HCl). Delivery address: Institut za rudarstvo i metalurgiju u Boru, Zeleni bulevar 35, 19210 Bor
4 | NANO spray dryer |  |  | 1 | 0 | 0
| 9051 | Laboratory instrument to produce particles with sizes in the nano range and small sample quantities (min. 200 miligrams)Using piazoelectric atomizing technology. Intended for work with aqueous and organic solutions. Equiped with: Chillerr with pressure control and oxigen gas monitoring Instrument for inlet air conditioning or condensation of water in closed loop configuration Electrical comunication with spray drying unit and solvent resistant tubings Electrostatic particle collector for particle separation Additional equipment for compressed gas:Unit for removal of oil from compresed air Air compressor, 5-7 bar final pressure, max 90 L/min Technical features: Evaporation capacity: 0,15 L/h Sample volume: 1ml to 200ml Particle size: from min 300nm to min 5 µm Droplet size range: from 8 to 21μm Nozzle types: Spray head (piazo technology, ultrasonic) Spray caps (hole diameter): 4µm, 5,5µm and 7µm Drying gas flow: 80 -160 l/min Heating temperature control: accuracy +/- 1°C Inlet temperature: Max 220°C Heating capacity: Max. 1400W Voltage: 200/230 V, 50-60 Hz Power: Max 1500 W Mean residance time: 1- 4s Spraying gas: compressed gas or nitrogen Interface: USB II Cooling unit - min. outlet temperature down to -25°C Inlet air conditioning unit - min. outlet temperature +2°C Additional equipment for compressed gas: Unit for removal of oil from compresed air Air compressor, 5-7 bar final pressure, max 90 L/min The producer is obliged to provide: Installation Operator training at customer site, immediately after the installation and confirmation of equipment specification Instruction manual Local service support by producer certified service technicians Delivery address: Institut za prehrambene tehnologije u Novom Sadu, Bul. cara Lazara 1, 21000 Novi Sad
5 | Microfluidic Bioanalyzer System |  |  | 1 | 0 | 0
| 9260 | General specifications • The system includes a chip reader, vortex mixer with chip adapter, chip priming station, operating software, data analysis software, brand-name laptop personal computer, color printer. • Allows for determination of size, integrity and concentration of nucleic acids (DNA and RNA and Small RNAs); size, purity and concentration of proteins; flow fluorescence cytometry analysis on a single platform. • Detection: laser-induced fluorescence (LIF) detector, blue LED detector. • The electrode cartridge removable from the electrode assembly • The electrodes removable from the cartridge Assays available DNA: • Sizing and quantitation of dsDNA fragments at least in range from 25 to 12000 bp. • Quantitative range: at least in range 5pg/µl to 50ng/µl • Can process at least 12 samples of DNA in parallel • Sample volume: around 1 µl RNA: • Quality control of total, transport and micro RNA samples. • Quantitative range: at least in range 50 pg to 500 ng • Can process at least 12 samples of RNA in parallel • Sample volume: around 1 µl Protein: • Sizing and quantitation of protein samples at least in range from 5 to 250 kDa • Can process at least 10 samples in parallel • Sample volume: around 4-5 µl Cytometry: • Cell assay kit for analysis of fluorescence parameters of cell samples • Can process at least 6 samples in parallel • Sample volume around 10 µl Analysis software • Single software for all available assays • Automatic digital archival of all data • data review software • Alternative data display options (gel-like image, electropherogram, dot-plot, histogram, and tabular formats) • RNA integrity algorythm • Direct comparison of multiple samples and chips • Multiple exportable data formats (text, image) Delivery address: Medicinski fakultet u Novom Sadu, Hajduk Veljkova 3, Novi Sad
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F9: Insert CIF price
Note on D10: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F11: Insert CIF price
Note on D12: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F13: Insert CIF price
Note on D14: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on F15: Insert CIF price
Note on D16: FILL all white fields, write brand, trade name, model and all relevant trade details