## Aux-Currencies (hidden)
EUR | 1 | 1 | 1 | 1 | 1 | 1
USD
RSD

## Sheet1
LOT 39 | Bidder: |  | TOTAL DAP PRICE OF LOT No39: |  |  | 0
Delivery address: Fakultet za fizičku hemiju u Beogradu, Studentski Trg 12-16, 11000 Belgrade, Serbia |  |  |  | 0
|  |  | Total EXW price |  | 0 | Date:
Shipping, handling and insurance
Line item No. | Reference No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Fluorescence detector with compact pump for HPLC |  |  | 1 | 0 | 0
| 1377 | Light source: xenon lamp, 150 W Wavelength range: 200-900 nm Bandwidth: less than 25 nm Wavelength accuracy: better than 5 nm Wavelength repeatability: better than 0.5 nm Sensitivity: S/N, Raman line of water 650 or greater Detector: photomultiplier or photodiode Cell volume: less than 10 μl Analog output: 1 mV or 10 mV Inputs: Remote chart work, Auto zero RS232: yes And the next features: Simultaneous measurements at up to 4 independent wavelengths. The detector must be connectable with the existing Bischoff HPLC system and must include the following parts: – High pressure isocratic pump with keypad and display (semi-preparative, flow rate range from 0.000 to min. 20.000 and not more than 25.000 ml/min with 1% flow accuracy and 0.001 ml/min increments, maximum operating pressure 200 bars, with automatic purge, flow calibration and compressibility compensation. The pump will be used for post-column derivatisation and for semi-preparative separations. – injector (for analytical and semi-preparative applications) – software to enable full pump control, flow programming and multipump gradient modes. Data acquisition and processing of the signal from fluorescence detector. – column for PAH’s.
Note on C1: Insert legal name of the Bidder
Note on E1: Choose from combo box EUR/USD/RSD
Note on F4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on G4: Insert Date
Note on F7: Insert CIF price
Note on D8: FILL all white fields, write brand, trade name, model and all relevant trade details