## Chromatography equipment
LOT 1 | Bidder: |  |  | TOTAL DAP PRICE OF LOT No1: |  |  | 0
Delivery address: Faculty of Chemistry, Studentski trg 14-16, Belgrade, Serbia |  |  |  | Total EXW price |  |  | Date:
Shipping, handling and insurance
Line item No. | Sub-item No. |  | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | HPLC semi-prep specification |  |  |  | 3 | 0 | 0
1.1 |  | Quaternary pump specification |  |  | 3
|  | Hydraulic system | Dual piston in series pump with servo-controlled variable stroke drive, floating pistons.
|  | Settable Flow range and operating pressure: | Settable Flow range: Set points 0.001 - 10 mL/min, in 0.001 mL/min increments.
Operating range up to 600 bar up to 5 mL/min
Operating range up to 200 bar up to 10 mL/min
|  | Flow presicion: | 0.07 % RSD and better
|  | Flow accuracy: | ±1% ili 10 μL/min or better
|  | Compressibility compensation | User selectable, based on mobile phase compressibility.
|  | Composition range: | 0 – 95% or 5 – 100% user selectable
|  | Composition precision | 0.2 % RSD
|  | Vacuum degasser | Number of channels: 4
Internal volume per channel: 1.5 mL
|  | Leak detection: | Leak sensor for leak detection, safe leak handling, leak output signal for shutdown of the pumping system.
|  | Diagnostic and GLP feature | Extensive diagnostics, error detection and reporting through LC software; continuous tracking of instrument usage in terms of seal wear and volume of pumped mobile phase with pre-defined and user settable limits and feedback messages. Electronic records of maintenance and errors.
1.2 |  | Autosampler specification |  |  | 3
|  | Injection range: | 0.1- 1500 μL in 0.1 μL increments
|  | Precision: | Better than 0.25 % RSD
|  | Operating pressure: | Up to 600 bara
|  | Injection cycle | 50 sec and better
|  | Cary over | 0.1% and better
|  | Leak detection: | Leak sensor for leak detection, safe leak handling, leak output signal for shutdown of the pumping system.
|  | Diagnostic and GLP feature | Extensive diagnostics, error detection and reporting through LC software; continuous tracking of instrument usage in terms of seal wear and volume of pumped mobile phase with pre-defined and user settable limits and feedback messages. Electronic records of maintenance and errors.
1.3 |  | Column compartment specification |  |  | 3
|  | Temperature range: | 10ºC below ambient up to 80 ºC and better
|  | Temperature stability: | ± 0.05 ºC and better
|  | Temperature accuracy: | ± 0.8 ºC and better
|  | Operating pressure: | 600 bar
|  | Column capacity: | At least three 300mm column
|  | Pre column heating and post column cooling | Two temperature zones with Peltier termosttating
|  | Heat up and cool down time: | 5 min from ambient to 40 °C
10 min from 40 °C to 20 °C
|  | Leak detection: | Leak sensor for leak detection, safe leak handling, leak output signal for shutdown of the pumping system.
|  | Diagnostic and GLP feature | Column identification module for GLP documentation of column type.
Extensive diagnostics, error detection and reporting through LC software; Electronic records of maintenance and errors.
1.4 |  | UVVIS detector specification |  |  | 3
|  | Detector type:: | 1024-element diode array
|  | Wavelength range | 190nm – 950nm
|  | Noise: | ± 0.7 x 10-5 AU and better
|  | Drift: | 0.9 x 10-3 AU/hr and better
|  | Wavelenght accuracy: | ± 1 nm, self-calibration with deuterium lines verification with holmium oxide filter
|  | Number of signals | 6 and better
|  | Slit width | Programmable: 1, 2, 4, 8, 16 nm
|  | Sampling rate | 80 Hz
|  | Flow cells | Analytical flow cell with 10mm path length operating pressure up to 100 bar.
Flow cell optimized for semi-prep application
|  | Leak detection: | Leak sensor for leak detection, safe leak handling, leak output signal for shutdown of the pumping system.
|  | Diagnostic and GLP feature | Extensive diagnostics, error detection and reporting through LC software; continuous tracking of instrument usage in terms of lamp usage with pre-defined and user settable limits and feedback messages. Electronic records of maintenance and errors.
1.5 |  | Refractive index detector specification |  |  | 3
|  | Detection type: | Deflection method
|  | Short term noise: | ±2.5 × 10 -9 RIU
|  | Drift | 200 x 10-9 RIU/h
|  | Refractive index range | 1.00–1.75
|  | Data Rate | Better than 35 Hz
|  | Temperature control: | From 5 °C above room temperature up to 55 °C
|  | Valves | Automatic purge and automatic solvent recycle.
|  | Leak detection: | Leak sensor for leak detection, safe leak handling, leak output signal for shutdown of the pumping system.
|  | Diagnostic and GLP feature | Extensive diagnostics, error detection and reporting through LC software; Electronic records of maintenance and errors.
1.6 |  | Fraction collector specification |  |  | 3
|  | Maximum system flow: | Up to 10 ml/min
|  | Delay volume: | Up to 50 μL
|  | Trigger modes: | Time slices and peak detection
|  | Fraction trays | 6 ml vials, 13ml test tube
|  | Trigger source | UVVIS or RID detector
|  | Leak detection: | Leak sensor for leak detection, safe leak handling, leak output signal for shutdown of the pumping system.
|  | Diagnostic and GLP feature | Extensive diagnostics, error detection and reporting through LC software; Electronic records of maintenance and errors.
1.7 |  |  |  |  | 3
Software for full instrument control, data analysis (UV spectra capability, UV library creation and searching) and report generation.
1.8 |  |  |  |  | 3
Brand name PC with monitor and laserjet printer according to the manufacturers specification and recommendation.
1.9 |  |  |  |  | 3
Systems must be configured with HPLC C18 analytical columns.
1.10 |  |  |  |  | 4x1
HPLC amino analytical and adequate semi prep column and analytical CHIRAL and adequate chiral semiprep column. Analytical column Ultron ER-OVM CHIRAL dimension 4.6 x 150 mm x 5 µm, semi-prep column Ultron ES-OVM CHIRAL, dimension 10 x 150 mm x 5 µm
2 | Gas Chromatograph Specification with FID and FPD and autosampler |  |  |  | 1 | 0 | 0
2.1 |  | Gas Chromatograph specification
|  | Temperature range: | From 4 °C above ambient up to 450 °C
|  | Oven ramp rates: | 2 °C/sec up to 110 °C
|  | Oven ramps/plateaus | At least 15 ramps and 16 plateaus
|  | Split/splitless inlet | Split/Splitless capillary GC inlet with bac-pressure control, pressure range up to 100 psi, pressure presicion 0.001 psi |  | 1
|  | Combined inlet | Should have possibility to work as split/splitless inlet and to be able to inject large sample volume . |  | 1
Mode of operation
-cold/hot split/splitless
-pulsed split/splitless
-solvent vent
-direct injection
Temperature range: from 10ºC above room temperature up to 450 ºC.
Number of programmed ramp: at least 9 ramp with heating rate of 900 ºC/min.
|  | Flame ionization detector | Flame ionization detector with electronic pneumatic control for all detector gases – Air, H2, and make-up gas. |  | 1
Linearity – up to 10e7
Minimum limit of detection < 1.5 pg C/s and better
Data rate: up to 500 Hz
|  | Flame photometric detector | Single wavelength detector
Minimum detection limit : less than 60ft P/s; less than 3.6pg S/s
Dynamic range: 10e3 for S, 10e4 for P
Data rate: up to 200 Hz
|  | Diagnostic and GLP feature | Extensive diagnostics, error detection and reporting through GC software; continuous tracking of instrument usage with pre-defined and user settable limits and feedback messages. Electronic records of maintenance and errors.
|  | Leak detection: | Leak detection, safety shut down if leak detected
Audio output signal when shutdown.
|  | Diagnostic and GLP feature | Extensive diagnostics, error detection and reporting through GC software; Electronic records of maintenance and errors.
2.2 |  | Autosampler specification
|  | Vial capacity | Up to 294 x 2ml viala; up to 96 10ml/20ml viala
|  | Injection mode | Liquid and headspace
|  | Syringe volume | Liquid from 1.2 µl - 500 µl
Headspace from 1ml – 5 ml
|  | Software control | Fully integrated into GC software
2.3
Software for full instrument control including autosampler, data analysis and report generation.
2.4 |  |  |  | `
Brand name PC with monitor and laserjet printer according to the manufacturers specification and recommendation.
3 | Gas Chromatograph Specification with FID, autoinjector and headspace sampler |  |  |  | 1 | 0 | 0
3.1 |  | Gas Chromatograph specification
|  | Temperature range: | From 4 °C above ambient up to 450 °C
|  | Oven ramp rates: | 2 °C/sec up to 110 °C
|  | Oven ramps/plateaus | At least 15 ramps and 16 plateaus
|  | Split/splitless inlet | Split/Splitless capillary GC inlet with bac-pressure control, pressure range up to 100 psi, pressure presicion 0.001 psi |  | 2
|  | Flame ionization detector | Flame ionization detector with electronic pneumatic control for all detector gases – Air, H2, and make-up gas. |  | 2
|  | Diagnostic and GLP feature | Extensive diagnostics, error detection and reporting through GC software; continuous tracking of instrument usage with pre-defined and user settable limits and feedback messages. Electronic records of maintenance and errors.
3.2 |  | Autoinjector specification
|  | Vial capacity | 15 x 2ml vials and better
|  | Minimum injection volume | 10 nl
|  | Maximum injection volume | Up to 250 µl
|  | Number of injection per vial | From 1-99
|  | Injection speed | Variable, depended from inlet type and mode of operation (split; splitless, on-column)
3.3 |  | Headspace sampler specification
|  | Vial size | 10ml, 20ml without adapter
|  | Number of extraction per vial | From 1 - 100
|  | Temperature zone | Oven: from 35ºC to 210 ºC
Sample loop: from 35ºC to 210 ºC
Transfer line: from 35ºC to 250 ºC
|  | Gas control of carrier gas and vial pressurization | Integrated
3.4
Software for full instrument control including autoinjector and headspace sampler, data analysis and report generation.
3.5
Brand name PC with monitor and laserjet printer according to the manufacturers specification and recommendation.
Note on C2: Insert legal name of the Bidder
Note on F2: Choose from combo box EUR/USD/RSD
Note on G4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on H4: Insert Date
Note on G7: Insert EXW price
Note on E8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G82: Insert EXW price
Note on E83: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G118: Insert EXW price
Note on E119: FILL all white fields, write brand, trade name, model and all relevant trade details

## Equipment for Material Char.
LOT 2 | Bidder: |  |  | TOTAL DAP PRICE OF LOT No2: |  |  | 0
Delivery address: Faculty of Chemistry, Studentski trg 14-16, Belgrade, Serbia Belgrade, Serbia |  |  |  | Total EXW price |  | 0 | Date
Shipping, handling and insurance
Line item No. | Sub-item No. | Technical Specification Requested |  | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Two in one ASTM Types I/II water |  |  |  | 1 | 0 | 0
1.1 |  | ASTM I
|  | conductivity | 0.055 μS/cm
|  | Resistance at 25°C | 18.2 MW×cm
|  | TOC Value | 1 – 5 ppb
|  | Bacterial content in CFU/ml | < 1
|  | Particle content | < 1
|  | Endotoxines in EU/ml | 0.005
1.2 |  | ASTM II
|  | Conductivity | 0,067 – 0,1 µS/cm
|  | Resistance at 25 °C | 15 – 10 MΩxcm
|  | Permeate performance at 15°C | 3 l/h
|  | Flow rate | 1.0 litre/min
Integrated ultrafilter, generously dimensioned for pyro­gene-free and nuclease-free water. The automatic self-cleaning gives years of service.
UV Oxidation 185/254 nm: UV oxi­dation for reducing the content of microorganisms and their metabo­lites.
Ultrapure water is dispensed through a standardly built-in exchangeble final sterile filter
Built-in 6 litre tank with pressure-/recirculation pump
Laboratory bench-top operated
Water Pretretman module(separate): Module 1 Pretreatment module with integrated reverse osmosis membrane; Module 2 Cartridge containing high-quality ultrapure resins with very long service intervals
2 | Autoclave, STANDARD, 75 Lit |  |  |  | 1 | 0 | 0
Vertical steam sterilizer with accelerated production of saturated steam through their own, separate electric steam generator. Compact and ergonomic design - high capacity and efficient utilization of space. A wide range of applications: microbiological culture, open liquids, glassware, instruments, non-porous material, waste. Operation control by PLC microprocessor installed on the ergonomic front panel mounted. Simple operation with the software. Large, clear color LCD screen with "soft touch" membrane. Display of temperature, pressure and temperature of showing graphics.
Minimum of 12 different, fully automated program to set free by the user.
Required characteristics: |  | • Sterilization chamber made of stainless steel 316 L, electro polished: • The supporting frames of stainless steel 304, the outer panels of stainless steel 304 and high resistance polyurethane. • Separate steam generator made of stainless steel 316 L with automatic pump to supply DEMI water. • Dual controls work - through control of pressure and temperature. • Built-quiet compressor (for instruments, valves and pneumatic sealing rubber) • Horizontal door with sliding open-close function (important) • Sealing ring silicone system that drives by air compressor (Steam or mechanics are not acceptable solutions) • 4 wheels for easy handling • Validation port • A flexible PT 100 temperature probes to control the internal temperature of the reference vessel • 4 levels of access to encrypted • Timer for programming the start of • Adjustable time for sterilization of 0-9999 min • Auto power-off with audible alarm • Disabled get started until the lid is closed • A separate fuel tank (tank) for supplying water DEMI • Safety devices to control water levels DEMI
|  | Temperature range: | 100 ° C - 138 ° C
|  | Maximum working pressure: | 2.5 bar g.
|  | The effective chamber dimensions (Ø x H) | 400 x 600 mm
|  | Volume: | 75 Lit
|  | External dimensions: | W 650 x D 650 x H 990 mm
|  | Power: | no more than 6.5 to 7 KW (energy saving)
|  | Electricity: | 400 V, three phase, + N + PE, 50 Hz
3 | Water Jacket laboratory CO2 Incubator, 184 Lit |  |  |  | 1 | 0 | 0
Constructions of chamber: water jacket - volume 43.5 Lit Water Jacket Technology – maximum thermal stability and quick, recovery are assured with triple wall construction.
Control of contamination: HEPA H14 filter class of circular shape with a system for easy replacement. HEPA air filtration systemontinuously filters the entire chamber volume every 60 seconds, up to air class 100.
temperature |  | Temperature range: | 5°C above ambient to 55°C
|  | Temperature Uniformity: | + / - 0.2 ° C to 37C
|  | Temperature control: | Temperature Sensor (TC)
CO2 |  | CO2 control: | better than + / - 0.1%
|  | CO2 range: | 0-20%
|  | Inlet pressure: | 1.0 bar
Temperature sensor
Humidity: ambient - 95% accurate with water Volume: 184 liters Lining the interior: Stainless steel 304, "mirror finished", rounded edges Lining the outside: cold rolled steel, painted Police: 4, stainless steel, perforated Outer door: Full Interior doors: 1, glass Internal dimensions: 66 cm (W) x 64 cm (D) x 100 cm (H) External dimensions: 54 cm (W) x 50 cm (D) x 68 cm (H)
Microprocessor control of: |  | • main switch "off / on" the front
• alarm-disorder temperature and CO2
• readable display with temperature and CO2
4 | Laminar flow cabinet for work with non-pathogenic microorganisms |  |  |  | 1 | 0 | 0
CE mark. Protection products and operator. For a 1 working place.
Designed with front air aspiration as 30% of the total volume of inside air with a speed of 0.40 m / s, with the remaining 70% that re-circulated through HEPA filter (as defined by European standard EN 12469)
|  | Outside dimensions: | 1312 x 1430 x 765 mm (wxhxd)
|  | Work zone: | 1220 x 640 x 580 mm (wxhxd)
|  | Work front opening: | 170-580 mm height
|  | Interior voulume: | 452 Lit
|  | Electr. port. | 230V single phase, 50 Hz
|  | Main filter: | 1 x H14 filter as per EN 1822 performance of 99.995% retention 0.3 μ particles, DOP / DEHS test.
|  | Interior air class: | ISO 5
|  | The safety net: | the main H14 filter made of "aluminum and anodised.
|  | The level of noise at work: | <58 dBA
|  | Brightness: | a fluorescent lamp in the uncontaminated area. (Minimum illuminance of 1000 Lux)
|  | Outer casing: | steel construction epoxy-polyester powder-painted
|  | Front window: | laminated safety glass with pneumatic lift.
|  | Working chamber: | stainless steel AISI 304L, with rounded edges
|  | Working surface: | stainless steel AISI 304L, multi-piece
|  | Maintenance. | Replacement filters and other activities on the front
|  | Volumes of air: | 335 m3 / h (exhaust), 1050 m3 / h (circulatory), 1385 m3 / h (total)
|  | Process Control: | The microprocessor guarantees motor activity and ensure optimal function.
Alarm: optical (red signal visible from the front panel work), and acoustic, active real-time indication of the type of alarms on alphanumeric display.
The existence of the display with representations: |  | Flow rate in m / s
Amount of air in M3 / h
Date and time
Date of last service
Ventilation system: |  | Fan centrifugal type with IP55 protection and automatic regulation.
Equipment: |  | Fluorescent Lamp 1 x 30W
Socket 4 and IP 55 (rear panel on the right)
UV lamp
Opening DEHS test.
Support - Steel Frame 795 mm, set up for working height of about 905 mm
5 | Dual-mode UV-Vis spectrophotometer with micro-volume and cuvette capability for nucleic acid and protein quantitation |  |  |  | 1 | 0 | 0
|  | Micro-mode: | Min. Sample Size: 0.5 µL
|  | Pathlength: | 1mm (auto-ranging to 0.05mm)
|  | Light Source: | Xenon flash
|  | Detector Type: | 2048-element linear silicon CCD array
|  | Wavelength Range: | 190 to 840nm
|  | Wavelength Accuracy: | 1nm
|  | Spectral Resolution: | < 1.8nm (FWHM at Hg 253.7nm)
|  | Absorbance Precision: | 0.002 A (1mm path)
|  | Absorbance Accuracy: | 2% (at 0.76 at 257nm)
|  | Absorbance Range: | 0.02 to 300 (10mm equivalent)
|  | Detection Limit: | 2ng/µL (dsDNA)
|  | Max. Concentration: | 15,000ng/µL (dsDNA)
|  | Measuring Time: | <5 sec.
|  | Sample Pedestal Material: | 303 stainless steel and quartz fiber
|  | Cuvette mode: | Heating: 37 ± 0.5°C
|  | Stirrer: | 150 – 850 rpm
|  | Path Length: | 10, 5, 2, 1 mm
|  | Absorbance Range: | 0.002 – 1.5
|  | Detection Limit: | 0.4 ng/µl (dsDNA)
|  | Maximum Concentration: | 750 ng/µl (dsDNA)
|  | Measurement Time: | < 3 seconds
Software Compatibility: Windows* XP (32-bit) with Service Pack (SP)2 or later, Windows Vista* (32 bit), and Windows 7 (32 bit and 64 bit)
6 | System of mercury porosimeters with following elements. |  |  |  | 1 | 0 | 0
Minimum features of the Porosimeters: |  | 1. Wet samples can be measured without change of materials moisture content. System need to have the vertical filling the degassing pressure and it can be set to the corresponding water vapour pressure (see also ISO 15901-1 regarding pre-treatment depending on materials nature).
2. The high pressure autoclaves comply with the European Directive for high pressure devices (97/23 CE –PED).
3. The 3-dimensional software must be available in modules.
4. The porosimeters must work independently or PC controlled. One PC can control four instruments
5. The instruments are equipped with serial interface RS232 for computer connection
6. The porosimeters must work without any liquid nitrogen or other gas
7. The porosimeters must work with dilatometers.
8. The porosimeters must simply optimises automatically the measurement by the samples nature without any pre-knowledge of the operator.
9. Data saving in ASCII and Excel, data report with free choice of results.
Position 1 |  | Macropore - Sample Preparation Unit incl. Ultramacropore (Including Software)
Fully automatic porosimeter and filling station for determination of pore volume and pore size- distribution up to the pressure of 400 KPa. Automatic vertical Mercury filling of the dilatometer under continuous vacuum and the correction of hydrostatic pressure depending from the height of the mercury column. Automatic double run for de-agglomeration of powders and determination of the particle size-distribution calculated from the space between the particles of a powder. Calculation of the apparent density.
Measuring ranges: Pore diameter from 58 µm up to 3.8µm (depending on contact angle of sample) Particle diameter from 330 µm up to 15 µm Pore volume max. 2.000mm3 max. data points 2.500 points Accuracy better than 0.1% (at constant temperature) Hysteresis better than 0.1% (at constant temperature) Transducer Resolution 0.01 up to 0.1 Kpa and 0.1 up to 400 KPa RS232-port Up to four instruments can run parallel with one PC.
Position 2 |  | High Pressure Porosimeter
Fully automatic porosimeter for determination of pore volume and pore size distribution Up to a pressure of 400 MPa The automatic detection system of the penetration steps decreases the total measuring time without any lost of resolution down to 25 minutes. The 9 different speeds for analysis allow the correction of the
Measuring ranges: Pore diameter from 15µm up to 0.0036µm (depending on contact angle of sample) Particle diameter from 40µm up to 0.01µm Pore volume max. 2.000mm3 Resolution 0,1mm3 max. data points 2.500 points Accuracy better than 0.1% (at constant temperature) Hysteresis better than 0.1% (at constant temperature) Transducer Resolution 0.01 up to 100 Mpa and 0.1 up to 400 MPa USB port Up to four instruments can run parallel with one PC.
Position 3 |  | Microstructure- Data- Station including Interface
Consisting of: Datastation, Monitor 17" TFT and Printer Incl. all necessary cables and software to control all the Porosimeters and evaluate the data.
Position 4 |  | Consumables for two years
Warranty period: minimum 2 years
7 | Refrigereted centrifuge with interchangeable rotors |  |  |  | 1 | 0 | 0
Temperature control: from -10C to +40C, with motor overheating protection, “pre-cooling function” -maximum speed 15200 rpm -maximum rcf 25830 -fixed angle rotor for 6x100mL with adapters for 50 mL conical tubes –max 13000 rpm (18516rcf) -fixed angle rotor for 48x2 mL – max speed 15 200 rpm (25 314 rcf) -swinging bucket rotor 4 x 750 mL – max speed 4700 rpm (4816 rcf) with 4 adaptors for 50 mL conical tubes (28 x 50 mL) and 4 for 15 mL conical tubes ( 56 x 15 mL). -noise level at maximal speed < 61dB -Tehnical standards :IEC 61010-1, IEC 61010-202, IEC 61010-2-101, 230 V : EN 292, EN 61326, EN 55011B -Certificated :UL listed, CSA certified, CE marked, IVD compliant, Certified Biosafety -User-friendly LCD back-lighted interface
8 | UV/VIS with Peltier System |  |  |  | 1 | 0 | 0
|  | Wavelength Range: | 190 to 1100 nm
|  | Wavelength Accuracy: | ±0.3 nm
|  | Wavelength Repeatability: | <0.1 nm
|  | Spectral Bandwidth: | Variable 0.5, 1.0, 1.5, 2.0, 4.0nm
|  | Optical System: | Double-beam
|  | Photometric Range: | >46A
|  | Photometric Accuracy: | ±0.001A @ 1A, ±0.002A @ 2A, ±0.005A @ 3A
|  | Data Collection Rate: | 50 points per second data collection rate (20 ms resolution; single cell)
|  | Source: | Xenon lamp
|  | Detectors: | Two matched silicon photodiode detectors
|  | Monochromator: | Modified Ebert monochromator
|  | Noise (RMS): | <0.00015A @ 0A; <0.0002A @ 1A; <0.0004A @ 2A
|  | Stability: | <0.0005 A/hr
|  | Stray Light: | <0.02%T @ 220; <0.005%T @ 340nm; <1%T @ 198nm (EP)
|  | Baseline Flatness: | ±0.001A
|  | Scan Speed: | 1 - 3800 nm/min
|  | Display: | LCD; 153x115mm, VGA graphics (Local Control models only)
|  | Supplied Interface: | RS-232-C (computer); Printer (parallel, PCL 4 or greater); external keyboard; DVI for external display, external trigger
|  | Test and Data Storage: | Floppy Disk (Local Control models only);
non-volatile memory
|  | Languages: | English
|  | Power Requirements: | 100-240V, 50-60Hz
Smart 8-Cell Peltier System, 220V
~ Precisely control temperature and automate measurements of up to 8 cells ~ Increase laboratory throughput ~ Measure multi-cell kinetics reactions under precise temperature control ~ Automate measurement of standards and samples for QA/QC where temperature control is required Features: ~ Full control of temperature set point and stirring ~ Precision engineered cell transport system offers quiet operation and positioning accuracy
~ Peltier temperature control ensures the highest accuracy and Reproducibility ~ Precision cell holder mechanism securely holds standard cells, microcells and flow cells ~ Integrated stirring coil for each cuvette ensures reliable stirring performance
Specifications |  | ~ Temperature range: 0 to 100°C ~ Temperature accuracy: ±0.1°C ~ Temperature reproducibility: ±0.05 °C ~ Sample Stirring speed: 250-2000 rpm ~ Holds up to 8 cells standard 10 mm pathlength cells
Advance software with 21 CFR Part 11
9 | Homogenizer with metal probe for emulsifications (making emulsions) of small volumes up to 1mL and homogenization of plant and animal tissue. |  |  |  | 1 | 0 | 0
Controllable rotor speed between 5000 and 39000 rpm. Metal tip (probe) with 5 mm diameter Stand for fixing homogenizer
10 | System with Fluorescence |  |  |  | 1 | 0 | 0
System with fluorescence, phosphorescence and bio-and chemi-luminescence measurement modes and excitation, emission , constant wavelength synchronous and constant energ synchronous spectral scanning; and kinetic scans consist of two systems:
1.1 |  | Multimode reader 100-240V, 50/60Hz
Technical specifications Multimode reader User interface The instrument is under PC control and run on SkanIt Software , which controls all the instrument functions and provides data reduction as well as reporting functions. Computer interface RS-232C Light source Xenon flash lamp Detector Photomultiplier tube for fluorometry (PMT1) and luminometry (PMT1 and PMT2) and a photodiode for photometry Measurement types Fluorescence intensity, time-resolved fluorescence, fluorescence polarization, absorbance and luminescence Wavelength selection By filters: Exitation (12.5 mm), Emission filters ( 25.4 mm) and Absorbance filters (12.5 mm), Plate types 6 to 384-well plates Fluorometry Excitation wavelength range 200 – 1000 nm Emission wavelength range 360 – 800 nm Sensitivity Fluorescence intensity: < 2 fmol fluorescein/well, 384-well plate Time-resolved fluorescence: < 20 amol Europium/well, 384-well plate Precision Fluorescence polarization: < 10 mP 1 nM fluorescein, 96-well plate Dynamic range Fluorescence intensity: > 5 decades, 384-well plate Time-resolved fluorescence: > 5 decades, 384-well plate Measurement time Fluorescence intensity: 5 – 1000 ms Time-resolved fluorescence: 5 – 10000 ms Measurement speed Photometry Wavelength range 200 – 1000 nm Linearity 0 – 2.5 Abs (96-well plates) at 450 nm, ± 2% 0 – 2 Abs (384-well plates) at 450 nm, ± 2% Measurement range 0 – 4 Abs Absorbance resolution 0.001 Abs Measurement time 20 – 1000 ms Measurement speed Luminometry Wavelength selection High-sensitive mode: no filters Standard mode: by filters Emission wavelength range Standard mode: 360 – 850 nm High-sensitive mode: – Sensitivity High-sensitive mode: < 10 amol ATP/well, 384-well plate Standard mode: < 200 amol ATP/well, 384-well plate Dynamic range > 5 decades, 384-well plate Measurement time 10 – 10000 ms Measurement speed Incubator Temperature range From ambient + 4°C to 45°C Shaker Shaking method Linear shaking Shaking amplitude 1 – 10 mm Dispenser(s) Dispenser(s) Up to two optional dispensers Syringe size 500 μl (1000 μl on request) Dispensing volume 5 – 500 μl with 1 μl increments Dispensing accuracy Dispensing precision Avoid volume 800 μl
1.2 |  | Fluorescence Spectrometer
Instrument includes: Optical Design Czerny-Turner monochromators, concave gratings 20 cm focal length Magnesium Fluoride coated optics Beam Geometry Horizontal, 5.0 mm wide Minimum Sample Volume 0.5 mL (standard 10 mm cuvette) Excitation Grating 1200 lines/mm, 250 nm blazed Emission Grating 1200 lines/nm, 400 nm blazed Excitation and Emission Spectral Bandwidths 0.5, 1.0, 2.5, 5.0, 10, & 20 nm (fixed, selectable) Light Source 150 W Ozone Free Xenon Lamp Sensitivity (Raman band of water) > 4000:1 RMS > 1000:1 Peak-to-peak Detector R-928 Photomultiplier tube Silicon photodiode for reference Zero Order Selection Excitation and Emission Long-pass filters for excitation and emission 320, 435, 530 nm open for zero order, 4 optional, user-determined Wavelength Range 190–900 nm Accuracy ±0.5 nm Repeatability ±0.2 nm Slew Speed 20,000 nm/min Scan Speed 1–6,000 nm/min Minimum Data Interval 0.1 nm Connectivity RS-232 with next items: Thermostatted 4-Position Fluorescence Cell Holder Fluorescence Microcell Holder Magnetic Stirrer (Automatic Type) FLR Cell 10x10 mm Quartz with stopper FLR Cell, 4MM wide Semi-Micro with stopper FLR Cell 225uL Micro Square Cell Quartz Flr Micro Cell Holder Adptr 3.30
Note on C2: Insert legal name of the Bidder
Note on F2: Choose from combo box EUR/USD/RSD
Note on G4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on H4: Insert date
Note on G7: Insert EXW price
Note on E8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G27: Insert EXW price
Note on E28: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G53: Insert EXW price
Note on E54: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G77: Insert EXW price
Note on E78: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G119: Insert EXW price
Note on E120: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G145: Insert EXW price
Note on E146: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G182: Insert EXW price
Note on E183: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G198: Insert EXW price
Note on E199: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G234: Insert EXW price
Note on E240: FILL all white fields, write brand, trade name, model and all relevant trade details

## 2-D Electrophoresys
LOT 3 | Bidder: |  | TOTAL DAP PRICE OF LOT No3: |  |  | 0
Delivery address: Faculty of Chemistry, Studentski trg 14-16, Belgrade, Serbia Belgrade, Serbia |  | Faculty of Chemistry, Studentski trg 14-16, Belgrade, Serbia | Total EXW price |  | 0 | Date
Shipping, handling and insurance
Line item No. | Sub-item No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | 1st dimension: IEF set with start-up kit |  |  | 1 | 0 | 0
"Integrated system for isoelectrophocusing with gold plated copper cooling plate complete with accessories and start-up set of IPG strips Line voltage: 115/230 VAC 50-60Hz Operating temperature: 15°C to 35°C Relative humidity : 0-70% Power consumption: 100W Output voltage: 0-10 000 V, resolution: 10 V Output current: 0-1.5 mA, resolution: 1 uA Maximum output power: 12 W Electrode area with cooling: Gold-plated copper IPG strip holders sizes: 7, 11, 13, 18, and 24 cm Strip holder capacity: 12 individual or manifold type carrier PC/software control: RS 232C/USB interface for PC control (incl. cable) with software Accessories: Manifold Light, equilibration set and box, paper bridges, electrode strips, cover fluid, 13cm ceramic strip holders 2. Start-up set consisting of 13 cm IPG dry strips: • pH 3-10 NL*, 5 packs • pH 3.5-6.NL, 1 pack • pH 5.3-6.5, 1 pack • pH 6.2-7.5, 1 pack • pH 7-11 NL, 1 pack"
2 | 2nd dimesnion: SDS-PAGE 18x24 cm set |  |  | 1 | 0 | 0
"Dual Cooled Vertical Unit for various electrophoresis, especially SDS-PAGE, running up to 4 slab gels of 18x24, 18x16 and/or 18x8cm size, including dual gel caster, heat exchanger, set of glass plates, 1.5 mm teflon spacers and 15-well combs, with set of additional accessories and start-up kit: Accessories: LF Glass plates 18x24cm set of 2, LF Glass Plates 18x16cm set of 2, Clamp and cam kit for 16 cm gels, 2 sets of 1.0 mm spacers and 2 each of 10-, 15-well 1.0 mm and 1.5 mm preparative combs. Start-up kit for fluorescnece blotting: ECF Western Blotting Reagent Pack 2 kits and 3 ECF Substrate kits."
3 | Laser scanning imager |  |  | 1 | 0 | 0
"Laser scanning variable mode imager with 1-D and 2-D analysis software including PC working station and accessories: Scanning plate size: 24 x 40 cm, scanning speed 2.5 min; Detection Modes: Fluorescence, phosphorimaging, chemifluorescence, filmless autoradiography and digitization (for visible stains) Optoelectronics: Excitation with 4 integrated lasers of 473nm, 532 nm, 635 nm and 650 nm with set of 4 emission filters; Range: up to 5 orders of magnitude; Detection: PMT Resolution: 25, 50, 100, and 200um. Accessorie: stage set for various types of samples 1-D and 2-D recording and analysis software (with integrated Melanie algorithm) including PC working station. "
4 | Graphite transfer system |  |  | 1 | 0 | 0
Graphite based 20x25 cm transfer system with 600 V power supply (as part of Multiphor-II kit family)
5 | Chromatography protein analysis and purification system |  |  | 1 | 0 | 0
"Chromatography system (2 pcs) for proteins, peptides, oligos and nucleic acids purification and process development system with PC working station, recording and commanding software, fraction collector, additional accessories and set of colums Gradient pump with 4 pistons; Flow: 0.001 - 10 ml/min; Upper pressure limit: 250 bar UV/VIS detector: Light Source: High Energy Xenon Flash Lamp, Wavelength Range: 190-700 nm, detection at 3 wavelenghts simultaneously Flow cell: analytical 8 ul cell with 10 mm optical path Fraction collector (2 pcs): carusel type for 95 standard type of tubes; Flow rate: 100 ml/min; microproccesor controlled multimodal automatic system PC working station with recording and commanding software for purification and process development use (2 pcs) Accessories (1 each): • short and long column holders, short column clamp, on-line filters, sample loops 10, 100 and 1000 ul Set of dedicated start-up columns as part of installation kit
6 | Flow Cytometer |  |  | 1 | 0 | 0
"- Automated flow cytometer with two lasers, one of 488 nm that detects 3 fluorescent colors and secondary red laser of 635 nm for detection of forth fluorescent color - Spatially separated lasers without spectral overlap, maximum sensitivity and possibility of fluorochrome choice which can be detected simultaneously - Choice of one out of three flow rates with maximum of 60 ul/min - Instrument for research application in immunophenotyping, absolute counting, counting of residual WBC, stem cell analysis and DNA analysis - Manual tube loading with peristaltic pump without tube loader for reducing sample carryover - Integrated software for automated compensation, adjustment and daily quality check with beads - Possibility of adding integrated module for sorting on flow in closed cuvette with module for cell concentrating to increase recovery after sorting - Possibility of adding integrated tube loader for automated sampling with possibility of choosing between manual and automated sampling - Possibility of adding integrated plate loader"
Note on C2: Insert legal name of the Bidder
Note on F2: Choose from combo box EUR/USD/RSD
Note on G4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on H4: Insert date
Note on G7: Insert EXW price
Note on E8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G15: Insert EXW price
Note on E16: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G23: Insert EXW price
Note on E24: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G31: Insert EXW price
Note on E32: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G37: Insert EXW price
Note on E38: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G45: Insert EXW price
Note on E46: FILL all white fields, write brand, trade name, model and all relevant trade details

## Preparative Chromatography
LOT 4 | Bidder: |  | TOTAL DAP PRICE OF LOT No4: |  |  | 0
Delivery address: Faculty of Chemistry, Studentski trg 14-16, Belgrade, Serbia Belgrade, Serbia |  |  | Total EXW price |  | 0 | Date
Shipping, handling and insurance
Line item No. | Sub-item No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | Gradient system for preparative chromatography |  |  | 1 | 0 | 0
1.1 |  | Piston pump |  | 2
Chemically resistant 3-piston pump module, constant, puls-free flow from min 3 ml/min to min 250ml/min or more. Maximum working pressure 10bar/145 Psi
1.2 |  | Pump controller for isocratic and gradient separations |  | 1
With an level sensor, overpressure sensor and mixing chamber
Possibility of adjustment of flow rate
LCD display for process parameters monitoring
2 Inputs/Outputs for solvent, with the possibility of solvent mixing and level sensors
1.3 |  | Injection unit for sample injection |  | 1
1.3.1 |  | Injection volume up to 5ml - 1 pc |  | 1
1.3.2 |  | Unit with 4-way valve with the possibility of loading volumes up to 20ml |  | 1
1.3.3 |  | Unit with 6-way valve with the possibility of loading volumes up to 20 ml |  | 1
1.4 |  | Fraction collector |  | 1
Collection of separated substances according to time, volume or peak
Possibility of work with the signals from 2 detectors
RS232 interface for data transfer to a PC.
Possibility of height adjustment depending on the height of rack and glass tubes
Run capacity: min 240 glass tubes or 12l
Has to be equipped with the valve for waste diverting and collection
Tube racks and glass tubes - 4 racks and 200 glass tubes
1.5 |  | UV VIS Detector |  | 1
Diode array detector
Wavelength: 200nm - 840nm
Accuracy of adjustment: ±1nm
Reproducibility: ±0,5nm
Simultaneous monitoring of 4 wavelengths
Light source: Deuterium and Halogen lamp
1.6 |  | RI Detector |  | 1
Analogue output 0-1V
LED light source, wavelength 950nm ± 20
Flow cell capacity min 22μl
Angle of deflection 5°
Max pressure 0,02MPa
Max detector flow 150 ml/min
Digital display for parameters monitoring
1.7 |  | Software for data acquisition |  | 1
Automated recording of all separation parameters with the possibility of forming reports
1.8 |  | Vacuum pump with Woulff bottle and manual vacuum contoler with needle\ |  | 1
Chemically resistant PTFE diaphragm pump with integrated gas ballast Final vacuum 10mbar Flow rate 1,7 m³/h Two membranes Maintenance free, without oil
1.9 |  | Columns, pre-columns and column filling sets:
Filling with silica gel |  | PP columns filled with silica gel - 10 pcs |  | 10
Particle size: 40 - 63 μm
Particle pore size: 60A
Column dimensions: diameter 40 mm, length 75 mm.
Packing weight: 40 g
Sample weight: 1000 mg
1
Set of empty PP columns with corresponding adapters for automatic filling on Buchi Cartridger C 670 |  | Empty PP columns - 50 pcs |  | 50
Diameter 12 mm
Length 150 mm
Filling weight 8 g
Sample weight up to 400 mg
Empty PP columns - 50 pcs
|  | Diameter 40 mm |  | 50
Length 150 mm
Filling weight 90 g
Sample weight up to 5000 mg
Set of empty glass columns, pre-columns and equipment for filling with silica gel |  | Glass column - 1 pc |  | 1
Inner diameter: 26mm
Pressure: 0-20bar
Column length: 460mm
Plunger column
|  | Glass column - 1 pc |  | 1
Innder diameter: 49mm
Pressure: 0-20bar
Column length: 460mm
Plastic coated
|  | Glass column - 1 pc |  | 1
Inner diameter: 26mm
Pressure: 0-40bar
Column length: 460mm
Plastic coated
|  | Preparative glass column - 1pc |  | 1
Prep elut Glass Column
Column for loading dry or barely soluble samples
Volume: 0-18ml
Pressure range up to 50 bar
|  | Pre-column - 1pc |  | 1
For columns diameter 15, 26, 36 and 49mm
Maximum allowed pressure 0-50bar
|  | Set for glass columns filling with silica gel by using compressed gas |  | 1
25-200μ Silica gel
For glass columns of inner diameter 15-49mm
|  | Set for wet filling and conditioning of glass columns with silica gel |  | 1
Particle size - less than 25μm
For column diameters 15mm or 26mm
1.1 |  | Column adapter |  | 1
for connection of columns into Biotage preparation system
Sytem has to have: Main voltage 220V and CE Certificates
The system should be delivered with all necessary parts and adapters in order to be ready for operation immediately
after delivery and installation.
All system parts should be compatible.
The supplier is obliged to provide:
Installation
Training at customer's site, immediately after installation and confirmation of equipment specification
Local service support provided by service engineers certified in the manufacturer training centers
2 | Gradient system for preparative chromatography |  |  | 1 | 0 | 0
1.1 |  | Piston pump |  | 2
Chemically resistant 3-piston pump module, constant, puls-free flow from min 3 ml/min to min 250ml/min or more. Maximum working pressure 10bar/145 Psi
1.2 |  | Pump controller for isocratic and gradient separations |  | 1
With an level sensor, overpressure sensor and mixing chamber
Possibility of adjustment of flow rate
LCD display for process parameters monitoring
2 Inputs/Outputs for solvent, with the possibility of solvent mixing and level sensors
1.3 |  | Fraction collector |  | 1
Collection of separated substances according to time, volume or peak
Possibility of work with the signals from 2 detectors
RS232 interface for data transfer to a PC.
Possibility of height adjustment depending on the height of rack and glass tubes
Run capacity: min 240 glass tubes or 12l
1.4 |  | Columns and adapters
Set of corresponding adapters for empty PP cartridges for automatic filling on Buchi Cartridger C 670
PP columns diameter is 12mm
Preparative elution adapter for loading dry or barely soluble sample, volume from 8 - 17ml, to be used withempty PP cartridges
50 pcs of corresponding frits for cartridges of 12mm diameter
Sytem has to have: Main voltage 220V and CE Certificates
The system should be delivered with all necessary parts and adapters in order to be ready for operation immediatelyafter delivery and installation.
All system parts should be compatible.
Note on C2: Insert legal name of the Bidder
Note on F2: Choose from combo box EUR/USD/RSD
Note on G4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on H4: Insert date
Note on G7: Insert EXW price
Note on E8: FILL all white fields, write brand, trade name, model and all relevant trade details
Note on G119: Insert EXW price
Note on E120: FILL all white fields, write brand, trade name, model and all relevant trade details

## System for mass desorption
LOT 5 | Bidder: |  | TOTAL DAP PRICE OF LOT No5: |  |  | 0
Delivery address: Faculty of Chemistry, Studentski trg 14-16, Belgrade, Serbia Belgrade, Serbia |  |  | Total EXW price |  | 0 | Date
Shipping, handling and insurance
Line item No. | Sub-item No. | Technical Specification Requested | Technical Specification Offered | QTY | EXW Unit price | Total Price per line item
1 | System for laser desorption and mass analysis |  |  | 1 | 0 | 0
The System for laser desorption and mass analysis consists of three connected devices:
i) A nitrogen laser that operates at 337nm; laser should produce 3ns duration pulses; the laser rate can be set to Default (3Hz) or Optimized; maximum possible rate 20Hz; compatibility with Voyager DE Pro mass spectrometer is required. Catalogue number: V400049
ii) and iii) Accompanying 3KV reflector detector system with power supply (PSU ASSY 3KV ELITE DETECTOR V725008 and CNTR MR DE-PRO, V700759). Compatibility with Voyager DE Pro is required.
Note on C2: Insert legal name of the Bidder
Note on F2: Choose from combo box EUR/USD/RSD
Note on G4: Include shipping, handling, insurance and all other costs to reach DAP price
Note on H4: Insert date
Note on G7: Insert EXW price
Note on E8: FIILL all white fields, write brand, trade name, model and all relevant trade details

## Aux-Currencies (hidden)
EUR | 2 | 1 | 1 | 1 | 1 | 1
USD
RSD