## Lot1 - Ele. Equ. and Tools
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 1: Electronics Workshop Equipment & Tools
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | Plastic sorting box - small | Stackable Plastic construction External Height: 50-100mm External Width: 100-200mm External Depth: 100-200mm | 100
2 | Plastic sorting box - medium | Stackable Plastic construction External Height: 200-300mm External Width: 150-200mm External Depth: 300-400mm | 75
3 | Plastic sorting box - large | Stackable Plastic construction External Height: > 200mm External Width: > 150mm External Depth: > 350mm | 50
4 | Solder wire 0.7mm | Solder wire 0.7mm Lead free External Diameter:0.7mm Material Composition: Lead Free, 99.3% Tin, 0.7% Copper Alloy Melting Temperature:227°C Weight:0.5kg Bit Temperature:350°C Max Bit Temperature:370°C Min Bit Temperature:350°C Wire Diameter:0.7mm | 2
5 | Bottle of flux | Flux Type: Liquid Flux Applications: Soldering Dispensing Method: Bottle External Depth: 50-55 mm External Width: 50-55 mm Length: 80mm Weight: 100g | 2
6 | 50 Lens Tissue Booklets in a Closeable Box, 25 Sheets per Booklet | Extremely soft, premium-grade, 124 mm x 73 mm in packs of five protective booklets, with 25 sheets per booklet | 3
7 | Alligator clips test leads | Alligator test leads Cable Length:500-600mm Connector Type A:Alligator Clip Connector Type B:Alligator Clip Set of 20 cables in 4 or 5 different colours | 28
8 | Test lead | Connector Type A:4mm Stackable Connector Type B:4mm Stackable Colour: Red Cable Assembly Type: Test Cable Length:1m | 105
9 | Test lead | Connector Type A:4mm Stackable Connector Type B:4mm Stackable Colour: Black Cable Assembly Type: Test Cable Length:1m | 105
10 | Test lead | Connector Type A:4mm Stackable Connector Type B:4mm Stackable Colour: Blue Cable Assembly Type: Test Cable Length:1m | 70
11 | Test lead | Connector Type A:4mm Stackable Connector Type B:4mm Stackable Colour: Yellow Cable Assembly Type: Test Cable Length:1m | 70
12 | Test lead | Connector Type A:4mm Stackable Connector Type B:4mm Stackable Colour: Green Cable Assembly Type: Test Cable Length:1m | 70
13 | BNC - Alligator test lead | Lead Length:2m Connector Type A:BNC Plug 50ohm Connector Type B:2 x Crocodile Clips Approval Bodies:IEC1010 30V ac, 60V dc Voltage Rating AC:30V Voltage Rating DC:60V Cable Length:2m Cable Type: Coax 50R No. of Ways:2 | 36
14 | Oscilloscope probe | attenuation: 1:1 i 1:10 max. measurement frequency: >100MHz x1 input capacitance: <100pF x10 input capacitance: <20pF working voltage: >500Vdc Test Probe Type: Oscilloscope Test Probe Functions: Frequency For Use With: Oscilloscope Lead Length:>1.2m | 45
15 | Soldering station with temperature regulation | Soldering station with tip cleaner. Min. Temperature: <160 °C Max. Temperature: >400 °C Low abrasive brass shavings tip cleaner No temperature drop | 10
16 | Vacuum pump | Min Flow Rate: 30 l/min Frequency: 50 Hz Gas Type: Air | 1
17 | Geiger Counter | Geiger Counter Radiation Detector, handheld | 2
18 | Gamma-Ray Spectrometer | Handheld digital gamma-ray spectrometer/Scintillometer Scan mode and data storage Weather protection Bluetooth connectivity | 1
19 | Balance | Electronic digital single plate balance Weight capacity: 0-5000 g Accuracy: max. 2 g Weight sensitivity: 1g Display type: LCD, four digits minimum Battery powered, optional AC adapter | 2
20 | Balance | Electronic digital single plate balance Weight capacity: 0-200 g Accuracy: max. 0.5 g Weight sensitivity: 0.1 g Display type: LCD, four digits minimum Battery powered, optional AC adapter | 2
21 | Micrometer | Max Measuring Range:25mm Graduations:0.01mm Graduation:0.01mm Min Measuring Range:0mm Resolution:0.01mm | 4
22 | Digital calliper | Capacity: > 150 mm Resolution: 0.01 mm | 4
23 | Alligator clip | Connector Type: Alligator Clip Max Jaw Capacity: 6mm Current Rating: 6A Insulator Colour: Black Termination Type: 4mm Socket Contact Material: Brass Contact Plating: Tin Mounting Type: Cable | 105
24 | Alligator clip | Connector Type: Alligator Clip Max Jaw Capacity: 6mm Current Rating: 6A Insulator Colour: Red Termination Type: 4mm Socket Contact Material: Brass Contact Plating: Tin Mounting Type: Cable | 105
25 | Alligator clip | Connector Type: Alligator Clip Max Jaw Capacity: 6mm Current Rating: 6A Insulator Colour: Blue Termination Type: 4mm Socket Contact Material: Brass Contact Plating: Tin Mounting Type: Cable | 70
26 | Alligator clip | Connector Type: Alligator Clip Max Jaw Capacity: 6mm Current Rating: 6A Insulator Colour: Green Termination Type: 4mm Socket Contact Material: Brass Contact Plating: Tin Mounting Type: Cable | 70
27 | Alligator clip | Connector Type: Alligator Clip Max Jaw Capacity: 6mm Current Rating: 6A Insulator Colour: Yellow Termination Type: 4mm Socket Contact Material: Brass Contact Plating: Tin Mounting Type: Cable | 70
28 | Screwdriver set 6 pieces | Screwdriver set with 6 pieces Parallel: 2.5, 3.5, 5.5mm, Phillips: PH1, PH2, Voltage Tester | 10
29 | Electronic side cutter | Overall Length:120-130mm Cutter Type: Side Cut Type: Extra Full Flush Cutting Capacity:>1.5mm Copper Wire Cutting Capacity:>1.5mm Edge Finish / Profile: Extra Full Flush Head Type: Oval Jaw Length:10-15mm Length:120-130mm Soft Iron Cutting Capacity:>1.0mm | 10
30 | Snip nose electronic pliers | Handle Type:P V C Jaw Type:Spipe Nose Bent (Serrated Jaw) Length:140-150mm Jaw Length:30-35mm Pliers Style: Long Nose | 10
31 | Insulated tweezers | Insulated nonmagnetic tweezers Tweezers Type:ESD Overall Length:140-150mm Handle Material: Insulated coating | 20
32 | Handheld antistatic desoldering pump | Handheld desoldering pump Tip Length:15-25mm Tip Diameter:3mm Body Material: Carbon | 20
33 | Ball end hex L-key | Ball end hex L-Key set, 9-PC Kit Contents:9 Pieces 1.5, 2, 2.5, 3, 4, 5, 6, 8, 10mm Hex Keys Hex Key Type: Ball End Long Arm Body Material: Chrome Vanadium Steel Material: Chrome-vanadium steel | 2
34 | Battery powered screwdriver | Battery Capacity: min 1.3Ah Battery Type:Li-Ion Max Torque:3N-m No Load Speed:150-200rpm Steel Drilling Capacity:min 6mm Weight:<500g | 2
35 | Set of electronics pliers | Kit Contents:4 Pieces Straight Pattern VDE Telephone Pliers VDE Combination Pliers Heavy Duty VDE Side Cutter | 2
36 | Screwdriver set 20 pieces | Kit Contents:20 Pieces Flared: 4, 5.5, 6.5, 8mm, Parallel: 2.5, 3, 4, 5.5mm, Phillips: PH00 to PH3, PZ0 to PZ3 Flared: 4 x 100, 5.5 x 100, 6.5 x 30, 8 x150mm Parallel: 2.5 x 50, 3 x 100, 4 x 100, 4 x 150, 5.5 x 150mm Phillips: PH00 x 50, PH0 x 75, PH1 x 30, PH1 x 100, PH2 x 125, PH3 x 150mm Pozidriv: PZ0 x 75 PZ1 x 100, PZ2 x 30, PZ2 x 125, PZ3 x 150mm Screwdriver Type: Phillips, Pozi, Slotted | 4
37 | Cabinet organiser 40-50 drawers | Cabinet Style: Drawer Cabinet Material: Steel Frame Material: Steel External Width: 300-350 mm External Depth:140-160 mm Drawer Material: Impact proof plastic Height: 500-550 mm No. of Drawers: 40-50 With plastic dividers for drawers, 3 per drawer | 16
38 | Conductive foam | Conductive foam for Integrated circuits storage Foam Type :High Density Colour: Black Height: 5-10 mm Width: 300-400 mm Length: 300-400 mm | 50
39 | SMD soldering/desoldering station | Professional soldering station with dual soldering irons. Soldering iron 150W, temp range 150 – 430 oC Chip tool for soldering/desoldering SMD components (desoldering tweezers) 2 x 20 W 150 – 400 oC With LCD display, including stands for both soldering irons. | 1
40 | Oscilloscope passive probe | Test Probe Type: Oscilloscope Approval Category: 55V CATII / 300V CATII Attenuation: switchable 1:1 / 10:1 Frequency Measuring Ranges: 30 MHz - 250 MHz Max Capacitance: 80 pF Resistance: 1 Mohm | 4
41 | PIC development system | PIC development system for PIC microcontrollers. With the PIC16F887 microcontroller. The system includes a USB 2.0 programmer with programming software and peripheral modules such as COG, port expander, MENU and 4x4 keypads, with GSM and accelerometer boards | 2
42 | PIC development system | PIC development system for PIC microcontrollers. With the PIC24 and dspic30/33 microcontroller. The system includes a USB 2.0 programmer with programming software for dspic30/33 and pic24. | 3
43 | PIC development system | PIC development system for PIC microcontrollers. With the PIC16F887 microcontroller. The system includes a USB 2.0 programmer with programming software and peripheral modules such as COG, port expander, MENU and 4x4 keypads | 16
44 | LV 32MX v6 development system MCU card with PIC32MX460F512L. | LV 32MX v6 development system MCU card with PIC32MX460F512L. | 1
45 | Programmable System on Chip development kit | Development board for Cypress Psoc with: Accelerometer,Thermistor, Proximity Sensing, CapSense | 1
46 | XtremeDSP Development Kit – Virtex-5 SXT Edition | ISE® Design Suite: DSP Edition (device-locked to the Virtex-5 SX50T FPGA) Cables & Adapters, Xilinx Platform Cable USB, Universal power supply CompactFlash Virtex-5 SX50T FPGA (XC5VSX50TFFG1136), DDR2 SODIMM (256 MB) ZBT SRAM (1 MB), Linear Flash (32 MB) System ACE™ CF (CompactFlash) interface Platform Flash, SPI Flash JTAG Programming Interface External Clocking (2 Differential Pairs) USB (2) – Host and Peripheral PS/2 (2) – Keyboard, Mouse RJ-45 – 10/100/1000 Networking RS-232 (Male) – Serial port Audio In (2) – Line, Microphone Audio Out (2) – Line, Amp, SPDIF, Piezo Speaker Rotary Encoder Video Input, Video (DVI/VGA) Output Single-Ended and Differential I/O Expansion GPIO DIP Switch (8), LEDs (8), and Pushbuttons (5) MII, GMII, RGMII, and SGMII Ethernet PHY Interfaces PCI Express® Edge Connector (x1 Endpoint) GTP: SFP (1000Base-X), GTP: SMA (RX and TX Differential Pairs) GTP: SGMII, GTP: PCIe®, GTP: SATA (Dual Host Connections) GTP Clock Synthesis Chips, Header for 2nd Serial Port 2nd Platform Flash PROM (32 Mb) for large device Mictor Trace Port, BDM Debug Port, Soft Touch Port | 1
47 | Bench drill | Min Speed Range: < 550rpm Max Speed Range: > 2500rpm No. of Speeds: 5 Height: 500-550 mm Max Chuck Capacity: 13mm Power Output: 250W Rotational Speed: > 2500rpm | 1
48 | Handheld drill | Supply Voltage AC: 220-230V Input Power: 125W Min No Load Speed (RPM): < 15000 Max No Load Speed (RPM): > 30000 Weight: 500-600g | 1
49 | Hammer drill | Chuck Size: 13mm Steel Drilling Capacity: 10mm Wood Drilling Capacity: 20mm Concrete Drilling Capacity: 12mm Chuck Size: 13mm Input Power: 600 W Weight: < 2.0kg | 1
50 | Bench grinder with sander | Power Rating app: 250W Blade/Disc Diameter: 150mm Grinding wheel - coarse grit, Sanding belt | 1
51 | Handheld Grinder | Disc Diameter: 110 mm Power: 500W No Load Speed: 1000rpm | 1
52 | Circular saw | Power: 1000 W No-load Speed: 5000 rpm Max Cutting Depth for 0° Bevel Angle: > 50mm Max Cutting Depth for 45° Bevel Angle: > 35mm Weight: 4.0 kg Spindle lock | 1
53 | Combined lathe/mill/drill machine | COMBINED LATHE/MILL/DRILL MACHINE Swing: 16-1/2 Distance Between Centres: 31.500 Spindle Bore: 1-1/16 RPM min - max: 150 - 1300 Horsepower: 3/4 Mill Spindle Type: 3MT Spindle Nose: 4MT Morse Taper Headstock: 3MT Morse Taper Tailstock: 3MT Spindle to Column (Decimal Inch): 11.20 Crosslide Travel Length (Inch): 8 Longitudinal Travel (Inch): 29 Tailstock Quill Travel (Inch): 4.333 Range Inch Thread Size: 4-120 Range Metric Thread Size: 0.75-6mm | 1
54 | Air compressor | Air Coupling Type: Quick Action Air Pressure: > 5bar Capacity: > 20l Air Delivery: > 100 l/min Power: 1.5kW Noise: < 90dB Mobile | 1
55 | Jigsaw | Power: > 600W Speed: > 3000 spm Steel Capacity: > 5mm Wood Capacity: > 75mm | 1
56 | BNC to chinch cable | Coaxial BNC to Chinch cable 12m | 15
57 | Laser cutter | CO2 Power: 50Watt Working area: 300mm x 500mm Precision: 0.01mm , resolution: 0.1mm Engraving deep: 0-0.1mm (wood or acrylic plastic) LCD for motion control and power control Interface: USB port Cooling: Water Cool Support: BMP,GIF,PCX,TGA,PLT,CDR,DXF,AI Drivers: CorelDraw V12, AutoCAD | 1
58 | Weather Station | Weather Station with 24-Hr Fan Aspirated Radiation Shield and wireless connection with at least 250m range. integrated sensors for rain collector, temperature, humidity, anemometer, solar radiation, UV sensor, and solar panel. | 1
59 | Metal construction worktable, width=200cm | Metal construction worktable heavy duty h=70-80cm w=200cm d=100cm with min 2 drawers With electrical installations 10 schuco type 220V, 2 UTP | 2

## Lot2 - Physics Lab. Equ.
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 2: Physics Laboratory Equipment
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | Power supply unit | Power Supply Output Type: Variable No. of Outputs:4 Output 1 Voltage:0-30V variable Output 2 Voltage:0-30V variable Output 3, 4 fixed or variable 5, 15V Output Current 1, 2: 3A Output Current 3,4:2A Load Regulation:0.01% Approval Bodies:CE Output Type: Adjustable Power Supply Type: Bench Supply Voltage:220V AC | 22
2 | Digital oscilloscope | Digital desktop instrument Oscilloscope Channels:2 Scope Bandwidth:25MHz Rise Time:<15ns Sample Rate:250 MS/s 2mV/div ~ 5V/div, 3% 1ns/div ~ 10s/div 0.01% CH1 , CH2 , Line , EXT AUTO, NORMAL, SINGLE, TV, Edge, Pulse AC , DC , LF, HF, Noise LCD display min. 5” USB and cardslot supporting Image (BMP) waveform data (CSV) and setup (SET) | 20
3 | Function generator with Counter | Sine, triangle and square wave function generator and counter Function generator frequency range: 0.05Hz-3MHz Counter frequency: >100MHz 2mV-20V peak-peak from 50 Ohm or 600 Ohm output Sine wave distortion: <2% on all ranges Square wave rise/fall times: <100ns Amplitude modulation capability | 20
4 | Digital bench multimeter | Digital bench multimeter Number of digits: 5 Voltage Measuring DC: up to 1000V, resolution from 0.1mV Voltage Measuring AC: up to 1000V, resolution from 0.1mV Current Measuring DC: up to 20A, resolution from 0.1uA Current Measuring AC: up to 20A, resolution from 0.1uA Resistance Measuring Range: up to 20M Ohm, resolution from 0.1 Ohm Capacitance Measuring Range: up to 50uF, resolution from 1pF Frequency Measuring Range: up to 200kHz | 22
5 | Digital handheld multimeter | Digital handheld multimeter Number of digits: 4 Voltage Measuring DC: up to 1000V, resolution from 0.1mV Voltage Measuring AC: up to 1000V, resolution from 0.1mV Current Measuring DC: up to 20A, resolution from 0.1uA Current Measuring AC: up to 20A, resolution from 0.1uA Resistance Measuring Range: up to 20MOhm, resolution from 0.1Ohm DMM Response Type: True RMS Battery Lifetime:400h | 24
6 | Rheostat | Track Resistance: 560 Ohm Track Taper: Linear Resistance Tolerance: ± 10% Resistor Element Type: Wirewound Termination Type: 4mm plug Max Current: 1A Travel: > 300mm | 2
7 | Rheostat | Track Resistance: 1000 Ohm Track Taper: Linear Resistance Tolerance: ± 10% Resistor Element Type: Wirewound Termination Type: 4mm plug Max Current: 1A Travel: > 300mm | 2
8 | Rheostat | Track Resistance: 4700 Ohm Track Taper: Linear Resistance Tolerance: ± 10% Resistor Element Type: Wirewound Termination Type: 4mm plug Max Current: 0.5A Travel: > 300mm | 2
9 | Capacitor decade box | Capacitor decade box Cap. Range: 100 pF - 10 uF Up to 250V | 2
10 | Digital Thermometer | Digital thermometer Pen Type Temperature Measuring Ranges °C:-50 to +150°C | 5
11 | Spectrum analyzer | Spectrum Analyzer Type: Bench Frequency Measuring Ranges: 100 kHz - 3GHz Accuracy: 5% / 10%, 1ppm/a (aging) Connector Type: BNC (Trigger/reference input) Display Reading: TFT Colour VGA Display, DVI output Input Impedance: 50 Ohm Interface Type: USB, IEEE-488, Ethernet/USB Max Load Voltage DC: >75V Memory Size: 1 memory curv Output Waveform: Tracking generator -20 dBm - 0 dBm (1dB steps) Range: 20ms - 1000s sweep Resolution: 1 Hz Safety Category: EN61010-1 Variable Gain: -114dBm / +20dBm | 1
12 | Signal Generator | Signal Generator Type: RF Bandwidth: 3000MHz Modulation Type: AM, FM, FSK, PSK and Pulse Accuracy: ± 5% FM, FSK,PSK, PhM and ± 4% AM Output Voltage: 2V Sweep Rate Range: 1MHz to 3000MHz Delay: 100 ns Impedance: 50ohm Output: -144 to +13dBm Total Harmonic Distortion: 3% | 1
13 | Oscilloscope | Oscilloscope Type: Digital benchtop instrument Oscilloscope Channels: 4 Bandwidth: 60MHz 2 mV/div to 10 V/div 2 mV/div to 5 mV/div: ± 4.0% full scale 5 nsec/div to 50 sec/div 5.5” or bigger colour lcd display QVGA Two usb ports, one for direct USB flash data download without PC CH1, CH2, CH3, CH4 and Ext trig inputs Built in FFT Memory Size: 20kpts No. of Channels: 4 Sample Rate: 2Gs/s Selectable 1 to 1,000 acquisition frames can be recorded, played back and stored in the scope memory or external USB memory 20 kpts half channel3, 10 kpts each channel | 1
14 | Logic Analyzer | Digital Logic Analyzer 18 ch PC Based With set of cables Frequency: 100MHz | 1
15 | Condenser microphone | Cardioid condenser design, class A FET electronics, 3 -micron, app 1” evaporated-gold mylar diaphragm, solid capsule, 10dB pad, switchable low-frequency roll-off filter | 1
16 | Microphone with switchable bass filters | Switchable bass roll-off filter protects against rumble, handling, pop and wind noise, on/off switch with battery test position, high output low noise, transformerless balanced output on 3-pin XLR connector | 1
17 | Super-cardiod microphone | Super-cardioid/lobar pick-up pattern Highly directional Low inherent self-noise High sensitivity Wide frequency response | 1
18 | Pattern cardiod microphone | Polar pattern- cardioid or hypercardioid Frequency range: 50 to 20,000 Hz Sensitivity app: 45 dBV Supply voltage: 9 to 52 V phantom power to DIN 45596 or internal 9 V battery (6F 22) Connector: 3-pin XLR | 1
19 | FireWire Audio Interface | 6 x 10 simultaneous analog/digital I/O 2 preamps Onboard DSP mixer 24-bit/192kHz 2-channel mic pre/2-channel A/D-D/A converter | 1
20 | HI-FI Stereo Studio Headphones | Circumaural Sensitivity: 90 dB/mW Frequency range: 16-2800 Hz Max input power 200 mW Additional velvet ear pads | 5
21 | Stereo Audio Mixer 8ch | Stereo Audio Mixer 8ch with USB | 1
22 | Colour Matching Cabinet | Colour matching cabinet with D65 International-standard artificial daylight (6500K) and TL84 (4000K) 18W App. 70X40x50cm | 1
23 | Professional eye tracker | Tracking Method: Infrared video. Dark pupil. Monocular or binocular options. Measurement principle The user can select between three methods: Pupil only, corneal reflection only, or both together. Accuracy Approximately 0.25° - 1.0° visual arc Spatial resolution* Approximately 0.15° visual arc | 1
24 | Luminance & Color Meters | Optical system 85mm f/2.8 lens; SLR viewing system; flare factor less than 1.5% Focusing distance 1014mm (40 in.) to infinity Angle of view 9° with 1° measurement area indication Luminance units: cd/m2 | 1
25 | Digital multimeter & thermometer | Infrared measurement range: -20 to 500°C General purpose thermocouple immersion probe (type K) measurement range: -200 to 1300°C Number of digits: 10 DC Range: to 1000 V DC Range: 10 A. AC Range: 750 Vac. AC Range: 10 A Resistance Range: up to 40 Megaohms. Capacitance Range: 400 uF. Soft carrying case Battery lifetime: 100 h | 3
26 | Desktop/portable Raman Spectrometer | Complete portable system with notebook-size computer with integrated software for full control of the system, data acquisition&storage, and spectrum database TE cooled CCD detector array Spectral range 200-3000 cm-1 Spectral resolution approx. 12 cm-1 Wavelength stability 1 cm-1 in 12 hours Photometric stability up to 4% in 12 hours 785 nm excitation wavelength, 5-300 mW laser power Fiber optics probe plus vial holder for liquids and powders Safety googles | 1

## Lot3 - Physics Teaching Set
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 3: Physics Teaching and Demo Sets
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | Robotics educational kit | Lego Mindstorms nxt 2.0 or equivalent system set including: Compass Sensor Sensor for LEGO Mindstorms NXT 2 x Gyroscopic Sensor for LEGO Mindstorms NXT 2 x Accelerometer Sensor for LEGO Mindstorms NXT 4 x Interactive Servo Motor for LEGO Mindstorms NXT Connector Cables for LEGO Mindstorms NXT Recharger for LEGO Mindstorms NXT Rechargeable battery | 2
2 | Fuel cell experiment kit | Didactic set for demonstration of hydrogen fuel cell energy process Include solar cell and electrolysis system for generating hydrogen and oxygen | 1
3 | Determining the gravitational constant | Didactic set: Gravitation torsion balance, for measuring with light source | 1
4 | Force as vector | Didactic set for showing composition and resolution of forces, with dynamometers, weights, hooks, and stands (or magnetic board based) | 1
5 | Block and tackle | Didactic setet of pulleys, weights, demonstration line, for constructing simple machines like block and tackle, spanish burton | 1
6 | Inclined plane | Didactic set: Adjustable angle inclined plane with trolley, for examination of composition of forces with dynamometers | 1
7 | Friction | Didactic set: Adjustable angle inclined plane with blocks, for examination of friction | 1
8 | One-dimensional motions on Fletcher’s trolley | Didactic set: Fletcher's trolley experiment with acquisition system for measuring trolley's speed. | 1
9 | Conservation of linear momentum | Didactic set: Two trolleys with ball bearinggs with track app. 1.5m, with two light barrier as movement detectors and measuring system. | 1
10 | Free fall | Didactic set for measuring free falling time with digital time measuring. | 1
11 | Conservation of angular momentum | Didactic set for examination of angular momentum conversation, with rotation model, and computerized measuring system. | 1
12 | Centrifugal force | Didactic set: Motor driven centrifugal force apparatus for measuring centrifugal force. | 1
13 | Moment of inertia | Didactic set for demonstration of moment of inertia with: torsion axle, set of cylinders and balls, and stands. | 1
14 | Simple pendulum | Didactic set: Simple pendulum for gravitational force measuring | 1
15 | Harmonic oscillations | Didactic set: Spring pendulum with magnetic sensor for measuring with data acquisition system. | 1
16 | Transversal and longitudinal waves | Didactic set: Helical spring and cord for demonstrations on longitudinal and transversal waves driven by dc motor | 1
17 | Propagation of water waves | Didactic set: Water waves generator with stroboscope equipped for showing refraction, reflection and interference. | 1
18 | Sound waves | Didactic set: Tuning fork with needle on the top for recording on carbon black. Including microphone with output for oscilloscope. | 1
19 | Oscillations of a string | Didactic set: Monochord with digital system for precision measurement of the oscillation frequency of a string is determined as a function of the string length and tensioning force. | 1
20 | Interference of ultrasonic waves | Didactic set: Three ultrasonic transducers 40kHz. | 1
21 | Barometric measurements | Didactic set: Liquid pressure gauge with U-tube manometer, glass vessel for liquid pressure gauge. | 1
22 | Bouyancy | Didactic set: Archimedes' cylinder with hydrostatic balance, set of weights, beakers, dynamometer 1.0 N | 1
23 | Surface tension | Didactic set: Set for measuring surface tension with break-away method. | 1
24 | Thermal expansion of solids | Didactic set: Al and Fe rods for measuring thermal expansion of solids, with erlenmeyer flask, alcohol burner, and stands with length measuring scales. Including apparatus for measuring thermal expansion of pipes with hot water flow, with controlled heater and pump | 1
25 | Gas laws | Didactic set: Including manual vacuum pump and gas thermometer for investigating the thermal behaviour (expansion, pressure increase) of gases, for introducing the absolute temperature scale, and for determining absolute zero. Thermometer consists of a capillary tube open at one end, in which a variable volume of gas is sealed with a drop of mercury. The open end is treated with gas-permeable frit and silica gel to prevent moisture from forming inside the tube. Includes a glass vessel for the water bath. | 1
26 | Basic experiments on electrostatics | Didactic set: Field electrometer, with glass and plastic rods, and stands | 1
27 | Displaying the equipotential lines of electric fields | Didactic set: Electrolytic tank for measuring the potential difference between the 0 V electrode and a steel needle immersed in the water. | 1
28 | Kirchhoff’s laws | Didactic set: Didactic set for demonstration of Kirchhoff's laws with resistors with connectors, cables, and protoboards. | 1
29 | Circuit with capacitor | Didactic set: Set of resistors and capacitors with connectors, cables and protoboards for determination of capacitor time constant. | 1
30 | Circuit with coil | Didactic set: Set of resistors and coils with connectors, cables and protoboards for determination of coil time constant. | 1
31 | Measuring-bridge circuits | Didactic set: Set of resistors, capacitors and coils with connectors, cables and protoboards for measurement bridges (Wien and Maxell) | 1
32 | Microwaves | Didactic set rof microwave examination, including: microwave gun oscillator with power supply, large horn antenna with set for microwave interference and diffraction. | 1
33 | Diode circuits | Didactic set: Set of diodes with connectors, cables and protoboards for demonstration of half and full wave rectifiers and voltage limiting with zener diode. | 1
34 | Transistors | Didactic set: Set of transistors, resistors and capacitors with connectors, cables and protoboards for examining circuits with transistors. | 1
35 | Reflection, refraction | Didactic set: Set for examination of reflection and refraction of light, with light source, mirrors and lenses, with stands or magnetic board based. | 1
36 | Laws of imaging | Didactic set: Set for measuring focal length of lens, with lens, stantds, and light source | 1
37 | Wind tunnel | Didactic set: School wind tunnel with: App. 300W suction and pressure fan Diffuser for inserting the suction and pressure fan Level bottom for experiments relating to the physics of flight and air resistance Bernoulli ramp (scaled) Sealing strip (scaled) Plastic cover (transparent) Rear wall (black) with orientation lines) Lift balance Scale for measuring the angle of incidence Pressure head after Prandtl | 1
38 | X ray thomography apparatus | Didactic set: X-ray apparatus: Dose rate at a distance of 10 cm: <1 µS/h Two independent safety circuits for doors, high voltage and emission current Automatic door locking: doors can be opened only, when no high voltages present i.e. no x-ray radiation can be generated. High voltage: 0 ... 35.0 kV (regulated DC voltage) Tube current: 0 ... 1.00 mA (independent regulated DC) Visible x-ray tube with molybdenum anode for characteristic short-wave radiation: Kα = 17.4 keV (71.0 pm), Kβ = 19.6 keV (63.1 pm) Including fluorescent screen for transillumination experiments: d = 15 cm and built-in rate meter including voltage supply for GM counter tube. With loudspeaker: as an acoustic ratemeter With goniometer with stepping-motor control Operating modes: manual control and automatic scan for sensor only, target only, 2:1 coupling Angular range: target unlimited (0° ... 360°) sensor -10° ... +170° Step width: 0.1° Exposure timer, gate time: 0.5 s ... 9999 s Bushings in the experiment chamber: high-voltage coaxial cable, BNC coaxial cable, empty channel for e.g. tubing, cable etc. Analog outputs: each proportional to target angle and to counting rate for chart recorder connection USB port for connecting a PC to control the x-ray apparatus, data recording and evaluation by the delivered software | 1
39 | Wave machine | Didactic set: For observation of propagation, reflection and superimposition of transverse waves; Consists of a chain of pendulums mounted on knife-edge bearings and turning about a common horizontal axis, with elastic coupling via rubber bands and brake on each pendulum including common break rail. All pendulums to be able to freeze simultaneously in any desired position. Min no. of pendulums: 20 Including driver for wave machine with range 0.1 – 1Hz Driver voltage range 5 – 12V | 1
40 | Data acquisition system with sensors | Data Acquisition system with sensors for didactic sets (items 5.9500 – 59542); 12bit resolution, up to 500 ksamples per second per input At least 5 analog inputs for sensors and voltage measurements USB connection with PC Including software Timer inputs, digital inputs With controlled output Including sensors for: Light barrier for detecting of passing – 4 pieces Rotation motion sensor – 1 piece Force sensor – 2 pieces Temperature sensor – 4 pieces | 2
41 | Hi voltage power supply | Hi voltage power supply, with regulation in range of 0-20kV, Max current up to 0.5mA Including hi voltage cables with 4mm banana jacks | 5
42 | Hi current power source | Dual output power supply with high max load Voltage range: 0-60V Max load at least: 20A | 5
43 | Transparent Tesla Coil with 12“ Toroid | Tesla Coil with toroid terminal top electrode 220 V, up to 4A electric supply Maximum output voltage 500 kV Transparent base box with visible inner electronics Earthed metal control panel Spark gap air cooled with electric fan Hardware in brass, primary coil copper tubing | 1
44 | NdFeB permanent magnets | NIB alloy N52, Disc form 2“ diameter X ½“ thick - 4 pcs NIB alloy N52, Square form 2“ X 2“ X ½“ - 4 pcs | 1
45 | Light Recording Pendulum | Floor-mounting steel construction Base diameter approx. 1000 mm Height approx. 2000 mm UV laser recording unit on photochromic surface Two-axes Blue and red recording plates | 1
46 | Floor Globe | Floor-standing stainless steel base (Galaxy or Columbus type) Globe diameter 60 cm Globe height 120 cm Rotating globe with political and physical map Switch-on internal globe illumination English language | 1
47 | Daylight Board | This electric board is a tool for demonstration of actual daylight parts of the globe including main associated information (GMT, Local time, duration of daylight/night, sunrise/sunset time, calendar, etc.). Wall-mounting aluminium frame Approx. dimensions: 90x60 cm Moving partially illuminated belt with world political map and legal time zones Clock displays with Greenwich Mean Time, Local Apparent time, Moment of sunrise/sunset English language Two spare lamps | 1
NOTE: For items 2.01-2.38 power supplies, function generators, oscilloscopes and data acquisition systems should not be included in the offer. Data acquisition systems (item 3) will be used for all listed items.

## Lot4 - Optics Lab. Equ.
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 4: Optics Laboratory Equipment
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | High speed camera | High Speed Colour Camera Capture speed up to 10000fps Expositure down to 100ns Min. Resolution 640x480 for max speed Recording length min 3 sec for max speed With local memory and expansion slot With AUX trigger and PC control 3 year warranty | 1
2 | Thermal imaging camera | Dual visual/IR image LCD Display Visual Camera Resolution: 2MPixel Microbolometer sensor resolution: 320x240 Image Frequency Refresh Rate: 60 Hz Min Temperature Range: -20°C to 500°C Accuracy: ± 1°C Min Operating Temperature: < -5°C Max Operating Temperature: > 40°C On Screen emissivity & transmissivity correction 3 year warranty | 1
3 | CCTV Camera | 1/2" CCD Sensor Composite Video Monochrome type Sensitivity: 0.0003 lux | 11
4 | Antivibrational optical table | Board: Optical table 3000mm x 1000mm x 310mm Weight: maximum 600 kg Metric hole structure 25 mm square matrix, M6 holes Surface: double-plated, corrosion resistant, black, magnetic steel Core single honey-comb Set of active antivibrational legs 600mm with self-levelling system + possibility of height adjustment min. 18mm Load capacity 2500kg Air compressor for active support frame and antivibration legs, max air pressure: 825 kPa | 1
5 | Antivibrational optical table | Board: Optical table 1250mm x 1250mm x 110mm Weight: maximum 200 kg Metric hole structure 25 mm square matrix, M6 holes Surface: double-plated, corrosion resistant, black, magnetic steel Core single honey-comb Active support frame with antivibrational legs 700mm with self-levelling system + possibility of height adjustment min. 18mm Load capacity 2500kg Air compressor for active support frame and antivibration legs, max air pressure: 825 kPa | 1
6 | Clamping fork | Clamping fork for positioning post assembly on optical bench | 5
7 | Set of rubber damping feet | Rubber damping feet, Set of 4, 1'' diameter, accepts M6 x 1.0 socket head cap screws | 10
8 | Sorbothane sheet 1/8'' | Sorbothane sheet, 12" x 12" x 1/8", hardness: 70 durometer, absorbs up to 95,5% of impact energy | 2
9 | Spectrometry education kit | Set of open-access elements for spectrometry training and demonstration in the spectral range of 200-880 nm: Wide spectrum small deuterium-tungsten light source with fiber-optic output Cuvette holder for square 1x1 cm cuvette and applicable various geometry of light source vs. spectrometer (for absorption spectrometry, turbidimetry, etc.) UV-VIS handheld spectrometer with USB connection output fiber-optics light cables with adapters | 1
10 | Tunable light source | Tunable light source range: 180nm -2000nm Changeable output windows for UV emitting control Arc xenon lamp (app. 75W) with lamp holder and tunable high intensity light source illumination. Additional (reserve) xenon lamp Monochromator with motor drive with computer software control Holographic grating 300–8000 nm +600 g/mm, 750 nm blaze. | 1
11 | CCD camera | Professional monochrome CCD camera USB connection Resolution 1600x1200, 12x9mm chip with approx. 7 micron square pixels Output Sensitivity 30 μV/e Thermoelectric cooling 16bit AD convector Antiblooming protection Speed download: 5 seconds for a full megapixel image Dark Current < 0.5 nA/cm2 3 year warranty | 1
12 | Camera lens adapter with filter wheel | 35mm Camera lens adapter cameras with filter wheel. For use with item no. 11 | 1
13 | Optical breadboard | Antivibrational breadboard, 750mm x 750mm x 60mm Metric hole structure 25 mm square matrix, M6 holes Surface: double-plated, corrosion resistant, black athermalized plated steel honeycomb core and internal damping magnetic stainless steel top and bottom plate Set of lifting handles | 3
14 | Spherical singlets lens kit | plano-/bi- convex/concave, lens set, ARC: 350-700 nm, 35 pc, diameter 1'': plano-convex f=25.4, 35, 50, 75, 100, 125, 150, 200, 250, 300, 400, 500; and bi-convex 750, and 1000; plano-concave lenses f = -30, -50, -75, and -100; bi-concave lenses f = -25, -50, and -75 | 1
15 | XYZ Translation Stage | For metric optical board Travel Range: 10 - 15 mm Vertical Load (Max) : 5 kg Horizontal Load (Max): 40 kg Orthogonality: <2 mrad Angular Deviation: <150 μrad Bearing Type: Ball on Hardened V-Grooves | 1
16 | XY linear translation with Z-axis rotation | For metric optical board Load: 4 kg, XY Travel Range: 10-15 mm Rotation range: 360 degree, bearings in V-Groove | 2
17 | 360 degrees rotational mount | For metric optical board Load: 4 kg (1.5 kg for fine adjustment) Vernier scale with 5 arc minute resolution Lockable. When locked +/-5 degrees fine adjustment possible with installed micrometer drive system Corrosion proof. Color: Black | 2
18 | Laser Quadrant Detector Sensor Head | For metric optical board Detection: 400 to 1050 nm, Measure X and Y coordinates of a beam’s position, two quadrant detector interfaces, 100 mW/cm2 damage threshold, Interface: 4-Ports, USB2.0 interface, 1000 Scans per Second, USB powered | 1
19 | Box with 10 absorptive Ø1" ND filters | OD=0.1, 0.2, 0.3, 0.4, 0.5, 0.6, 1.0, 2.0, 3.0, and 4.0 filters, Schott glass, housed in SM-compatible mounts, in convenient storage case | 1
20 | Power Meter with Sensor | For metric optical board Power meter: USB2.0 interface 16bit A/D convertor included graphical user interface Sensor: Wavelength range: 400 - 1100 nm optical power range: 500nW - 500 mW type: Si resolution: 10nW | 1
21 | Translation Stage with Differential Micrometer | For metric optical board Adjustment range app: 10 mm fine adjustment range: 250 μm fine adjustment graduation (per division): 0.5 μm max load: 40 kg, Lockable Coarse Adjustment | 2
22 | Motorized Linear Translation Stage with Controller | For metric optical board Travel Range 50mm (2”) Bidirectional Repeatability 1.6um Backlash <6um Velocity 3mm/s Vertical Load min 4kg Horizontal Load min 10kg Including servo motor controller with USB connection and power supply | 1
23 | Motorized Rotation Stage with Controller | For metric optical board Continuous 360° Rotation DC Servo Motor with Gear Head and High-Count Optical Encoder (at least 10,000 pulses/rev) Angular Readout in 2° Increments 5 arcmin Precision min 10 kg Vertical Load Including DC servo motor controller with USB connection and power supply | 1
24 | Protected silver mirrors | Pack of 10, Ø1", substrate material: fused silica, clear aperture: >90% of diameter, thickness tolerance: ±0.2 mm, laser damage threshold: 0.5 J/cm2 (10 ns pulse) | 1
25 | Cylindrical Laser Mount | For metric optical board Two axes of precision angular adjustment, Ø50 mm maximum and Ø15 mm minimum clamping diameter, ±3° pitch and yaw adjustment | 2
26 | Mounted Retroreflector | For metric optical board Diameter: 50 mm, L=40 mm, AR-Coated: 350-700 nm, clear aperture: 70% of area, surface flatness: λ/10 at 633 nm, Beam Deviation: 3 arcmin | 2
27 | Non-Polarizing Beamsplitter Cube | 25.4mm, AR-Coated, 50:50, 400-700 nm, Clear Aperture: >80% of Cube Face, Reflected Beam Deviation: 90° ± 5 arcmin, Transmitted Beam Deviation: <5 arcmin (Max), Surface Flatness: λ/10 at 633 nm, Splitter Ratio Tolerance: <15% Over Entire Wavelength Range | 2
28 | Non-Polarizing Beamsplitter Cube | 10mm, AR-Coated, 50:50, 400-700 nm, Clear Aperture: >80% of Cube Face, Reflected Beam Deviation: 90° ± 5 arcmin, Transmitted Beam Deviation: <5 arcmin (Max), Surface Flatness: λ/10 at 633 nm, Splitter Ratio Tolerance: <15% Over Entire Wavelength Range | 2
29 | M4 cap screw kit | M4 cap screw kit, over 1000 pieces, 6, 10, 12, 20, 25, 30, and 40 mm long socket head cap screws; nuts and washers | 2
30 | M6 cap screw kit | M6 cap screw kit, over 1000 pieces- 10, 12, 16, 20, 25, 30, 35, and 45 mm long socket head cap screws, M6 setscrews with 12 and 20mm lengths, nuts and washers | 2
31 | M4 metric screws set | M4 stainless steel screws; 4, 5, 6, 10, 12, 16, 20, and 25 mm long; 2.0 mm Hex Keys(10pc). | 2
32 | M6 metric screws set | M6 stainless steel screws: 6, 10, 12, 16, 20, 25, 30, and 35 mm long; 3 mm Hex Keys(10pc). | 2
33 | M3 metric cap screw set | 1000pc M3 cap screws: 5, 6, 8, and 10 mm long socket head cap screws; 6 mm long setscrews; nuts and washers; ten 1.5 and ten 2.5 mm hex wrenches | 2
34 | Metric hex key set | Two hex key organizers, sizes: 0.7, 0.9, 1.3, 1.5, 2, 2.5, 3, 4, and 5 mm(20pc each), and 6mm (10pc) | 1
35 | Polarizing beamsplitter cube | Size: 10mm: wavelength range: 420 - 680 nm, transmission efficiency: Tp > 90%, reflection efficiency: Rs > 99.5%, extinction ratio: Tp:Ts >1,000:1, clear aperture: >80% of dimension, transmitted beam deviation: <5 arcmin, surface flatness: >λ/10 @ 633 nm, surface quality: 40-20 scratch-dig | 2
36 | Polarizing beamsplitter cube | Size: 20mm: wavelength range: 420 - 680 nm, transmission efficiency: Tp > 90%, reflection efficiency: Rs > 99.5%, extinction ratio: Tp:Ts >1,000:1, clear aperture: >80% of dimension, transmitted beam deviation: <5 arcmin, surface flatness: >λ/10 @ 633 nm, surface quality: 40-20 scratch-dig | 1
37 | Compact Ø1.5" Post Mounting Clamp, Metric | M4 x 0.7 tapped mounting holes to be attached to a Ø1.5" post, providing the setting of the angle of the optic mount via the knurled knob, with four holes provided on the knob to allow a 2.0 mm hex wrench to be used as a torque wrench to provide final locking | 4
38 | Visible alignment disk | size: Ø1", IR alignment disk, working ranges: 400-640 nm and 800-1700 nm | 4
39 | Mounting post for lasers | Ø1.5",solid stainless steel, the ability to stack multiple mounting posts, length 350mm, metric hole | 3
40 | Mounting post for lasers | Ø1.5",solid stainless steel, the ability to stack multiple mounting posts, length 200 mm, metric holes | 3
41 | 15-Piece Balldriver and Hex Wrench Kit with Stand, Metric | Six balldrivers ranging from 1.5 mm to 5 mm + nine hex keys ranging from 0.7 mm to 5 mm, with stand, metric | 2
42 | Kinematic Prism Mount, Metric | Kinematic prism mount, metric, for aligning components that require pitch and yaw adjustment, platform with M4x0.7 tapped holes | 2
43 | Large clamping arm, M4 threads, metric | Large mounting surface with a multitude of various mounting holes, fits both 30mm and 60mm cage systems, metric | 3
44 | Small clamping arm, M4 threads, metric | Small mounting surface with a multitude of various mounting holes, fits both 30mm and 60mm cage systems, metric | 4
45 | Mini light with flexible arm, mounting clip, and magnetic base | White light source, a flexible arm, attaching via the clip fastener or the magnetic base, batteries AAA | 3
46 | Right angle bracket with two 1.035"-40 tapped clear apertures | Right-angle plates with threaded clear apertures, rigid mounting structures for direct use within the beam path, maximum rigidity with parallelism and perpendicularity to within 0.05 mm, metric | 2
47 | General purpose plate holder | Jaw range app 15mm thick | 2
48 | 4-way cuvette holder | Cuvette holder for square 1 cm cuvettes Equipped with four fiber optic couplings on each side with appropriate f/2 collimating lenses Built-in 1/4" filter slot Base plate with optic fiber holders | 1
49 | Bases and post holders set | Metric post holders with spring-loaded hex locking thumbscrew, Ø1/2", lengths: 1''(10pc), 1.5''(10pc), 2''(10pc), 3''(10pc), 4''(5pc), 5''(5pc); (with thumbscrews included); L shape general purpose table clamp (~20pc); mounting bases: 25.0 x 75.0 x 10.0 mm(10pc), 25.0 x 60 x 10.0 mm(10pc), 50.0 x 75.0 x 10.0 mm(10pc) | 1
50 | Mounting posts and accessories set | Metric stainless steel posts with removable M4 mounting stud on top and M6 tapped hole on other end, sizes(10pc each): 20mm, 40mm, 50mm, 75mm, 100mm, 150mm, 200mm; V-clamp for large cylindrical objects such as laser tubes, sizes(2pc each): Ø0.5 mm to Ø22 mm and Ø14 mm to Ø50 mm; small clamping arm(2pc); large clamping arm (4pc); right-angle clamp, facilitates the alignment of two Ø1/2" posts oriented at 90° with respect to one another(10pcs); swivel post clamp, allows two Ø1/2" posts to be set at any angular orientation within a plane and provides 360° continuous adjustment(3pcs); 90° angle bracket, flat, parallel, and perpendicular to within 0.05 mm, (M6) clearance slots, 50x50mm (5pcs); includes Ø0.5" posts, right-angle clamps, swivel clamps, and V-clamps | 1
51 | Optical Mounts Essentials Kit | Kinematic mirror mounts, Ø1" optics, brass bushings, pitch and yaw angular range of ±4° (10pcs); kinematic platform mounts(3pcs): 2''x2'', provides ±4° of pitch and yaw motion; kinematic platform mounts(2pcs): 3''x3'', provides ±3° of pitch and yaw motion; kinematic prism mount(2pcs): 1.5''x1''; precision right angle plate(2pcs): stackable filter holder, soft nylon coating on spring clips to protect filters(3pcs); fixed optic mount, 1'' (5pcs) 2'' (5pcs) with retaining rings; adjustable lens mount: Ø5 mm to Ø45 mm(3pcs); adjustable lens mount: Ø19 mm to Ø57 mm(3pcs); self-centering lens mount, Ø3.8 mm to Ø43 mm(2pcs); standard iris diaphragm, Ø12.0 mm max aperture(3pcs); rotation platform(2pcs); rotation mount for Ø1'' optics(2pcs); 45° mounting adapter(2pcs); 45° optic holder(2pcs); Metric | 1
52 | Pillar Post Essentials Kit | Includes: Ø1" Pillar Post Extensions, lengths 0.5'', 1'' , 2'', 3'', 4'' (5pc each) with M6 taps; pedestal pillar posts Ø1" lengths 0.5'', 1'' , 2'', 3'', 4'' (5pc each) with M6 taps; pillar post extension with a hole Ø0.5"', 1'' , 2'', 3'', 4'' (5pc each) with M6 taps;; small clamping forks Ø1.25" (10pc), Ø1.75" (10pc); pillar post c- clamp Ø1" (10pc); pillar post angle c-clam Ø1" (10pc); pillar post adjustable height c-clamp (5pc); 1" construction cube w/ counterbored clearance holes (5pc); 1" construction cube w/ 1/4"-20 tapped holes (5pc); Metric | 1
53 | Cage Assembly and Lens Tubes Essentials Kit | Ø2" lens tube, 0.5" bore depth, retaining ring included, (5pcs); Ø2" lens tube, 1" bore depth, retaining ring included, (5pcs); Ø2" lens tube, 2" bore depth, retaining ring included, (5pcs); 60 mm threaded cage plate, (3pcs); 30 mm to 60 mm cage plate adapter, (3pcs); z translation mount for 30mm cage systems, translation along optical axis, (2pcs); xy translation mounts for 30 mm cage systems with 100 tpi drives, (pcs); mounting adapter for Ø11 mm collimators, (2pcs); rod cross coupler, top and bottom pieces plus four 2-56 x 1/4 caphead screws (5pcs); set of 4 swivel couplers (5pcs); rod end swivel connector, set of four (5pcs); 30mm cage rotation mount for Ø1" optics, 360 degrees rotation (3pcs); 30 mm cage assembly, xy translating lens mount, compact design 6.5mm travel (3pcs); retaining ring for Ø2" lens tubes and mounts (2pcs); lens tube caps, internal thread (2pcs); lens tube caps, external thread (2pcs); Metric holes | 1
54 | Cage assembly and lens tubes essentials kit, metric | Includes: 10 Pack of 0.05" Hex Keys; Cage Assembly Rod, 1", 1.5", 2", 3", 4", 6", 8" Long, Ø6 mm; SM1 Threaded 30 mm Cage Plate, 0.35" Thick SM1-Threaded Cage Plate, 0.5" Thick Lens Tube, Ø1"- 0.50", 1'', 2'' Long Ø1" Rotating Adjustable Focusing Element- 1/2", 1'' Travel 4-Way Mounting Cage Cube Lens Tubes adapters Lens Tube, Ø1/2", 0.50" Long(Ø1.035"-40) Locking Ring, Ø1.25" Outer Diameter Slim Lens Tube Slip Ring Lever-Actuated Iris Diaphragm (Ø0.8 - Ø12 mm) (Ø1.035"-40) Coupler, Internal Threads (Ø1.035"-40) Coupler, External Threads Blank Cover Plate with Rubber O-Ring Tapped Cube Plate Rotatable Cage Cube Platform Rotatable Kinematic Cage Cube Platform Ø1" Cage Cube Optic Mount 30 mm Cage System Alignment Plate with Ø1 mm Hole Coarse ±1 mm XY Slip Plate Positioner Threaded Kinematic Mount for Ø1" Optics | 1
55 | Lens Mount and Adapters Essentials Kit | Contains 50 lens mounts, 50 adapter rings for smaller diameter optics, and two spanner wrenches for retaining ring installation, with M4 x 0.7 mounting holes, in a nine-drawer cabinet (includes: lens mount for Ø1/2" optics, lens mount for Ø18 mm optics, lens mount for Ø1" optics, lens mount for Ø30mm optics, lens mount for Ø1.5" optics, lens mount for Ø2" optics, lens mount for Ø75mm optics, adjustable spanner wrench, spanner wrench) Metric | 1
56 | Spanner wrench | Spanner wrench with clear aperture for installation and adjustment of retaining rings, lens cells, and aspheric lens adapters, diameter 1'' | 2
57 | Adjustable spanner wrench | Spanner adjustment ranges from 3 mm to 75mm, one-handed operation adjusts spanning range, replaceable steel blades Spare blade | 2
58 | Diode Pumped Solid State Laser | Compact diode pumped solid state laser Output wavelength 532 nm Power app: 1W Continuous wave (cw) Beam quality TEM00 (M2 < 1.1) Beam roundness: 1.1 Thermal tuning coefficient [GHz/K]: -6 Thermal tuning range [GHz]: 60 Thermal response bandwith [Hz]: 1 PZT tuning coefficient [MHz/V]: 2 PZT tuning range [MHz]: ±200 Emission spectrum: Single-frequency Spectral linewidth [kHz/100ms]: 1 Coherence length [km]: > 1 Frequency stability 1) [MHz/min]: 2 Intensity noise, 10 Hz to 2 MHz [% rms]: < 0.1 Including power supply 3 year warranty | 1
59 | Diode System Laser | Laser Class - CDRH IIIb Output Power (mW): 100 Wavelength (nm): 640 Beam Diameter, 1/e2 (mm): 1 Beam Divergence (mrad): max 1 Spatial Mode: TEM00 Mode Quality, M2 <1.5 Polarization >100:1 Linear Warm-Up Time (minutes) <5 Digital Modulation: Maximum Bandwidth: 150 MHz Rise or Fall Time (10% to 90%): <2ns Modulation Depth (Extinction Ratio): >1,000,000:1 at 0 Hz, >250:1 at 150 MHz Analog Modulation: Maximum Bandwidth: 350 kHz Rise or Fall Time (10% to 90%): <1μsec Modulation Depth (Extinction Ratio): >10,000:1 Pointing Stability after Warm Up (mrad/°C): <6 Power Supply 220V AC, 50Hz RoHS Compliant, CE Certified 3 year warranty | 1
60 | Blue Diode Laser Module | 405 nm, min 4 mW, Laser Diode Module, Class 3R In 1.5” laser cage with shutter and power switch Including power supply | 1
61 | Optical Trapping Educational Kit with Force Measurement Module | Complete Optical Trapping Educational Kit with Force Measurement Module (Metric) Including: - Laser Diode Source app. 300MW, less than 1000nm - Laser Diode controller and power supply - 3-Axis Sample Positioning Stage - CCD Camera with USB Interface for Video Imaging - Position-Sensing Detector Module Available - Force measurement modules - Beam expanders and objectives - Mounting tools and parts - Software for imaging and measuring | 1
62 | Filter Kit | 12.5mm Mounted Diameter Filter Kit, Range 400-700nm CWL minimum 40 Mounted Filters of different wavelengths 10nm FWHM | 1
63 | Lens tube, long 1", with 2 retaining rings, 1'' diameter | Lens tube, long 1", with 2 retaining rings, 1'' diameter | 2
64 | Lens tube, long 2", with 2 retaining rings, 1'' diameter | Lens tube, long 2", with 2 retaining rings, 1'' diameter | 2
65 | Fish-eye lens | Fish-eye lens, f=1.8 mm, f/1.4, for 1/2" CCTV cameras Manual Iris Fixed Focus | 4
66 | Fish-eye lens | Fish-eye lens, f=1.4 mm, f/1.4, for 1/2" CCTV cameras Manual Iris Fixed Focus | 1
67 | Camera lens | For 1/2" CCTV cameras 6.0mm f0.8 Video Type Lens Aspherical Auto iris | 9
68 | Security camera case | For Item No. 67 and 3 | 9
69 | Camera case allsky cameras | For Item No. 3 | 3
70 | BNC to chinch cable | Coaxial BNC to Chinch cable 6m | 15
71 | Transmission Gratings | Dimensions (mm) 50 x 50 Thickness (mm) 3 Nominal Dimensional Tolerance (mm) ±0.5 Thickness Tolerance (mm) ±0.5 Groove Density (grooves/mm) 600 | 2
72 | Camera stand | Camera stand Tripod height adjustable: 50-130cm | 2
73 | Adapter 220V AC / 12V DC 1A | Adapter 220V AC / 12V DC 1A, euro plug | 10
74 | Laser protection goggles | Fitover style (style code 33 or 60) Soft temples Top and sideshield protection; For green laser light (532 nm) EN207 standard compliance | 6
75 | Laser protection goggles | Fitover style (style code 33 or 60) Soft temples Top and sideshield protection; For red laser light (600-700 nm) EN207 standard compliance | 6
76 | Handheld XRF | X-ray tube Silicion Drift Detector Li-on batteries Waterproof rugged carrying case USB connection PC Software Soil Exploration: K, Ca, S, P, Cl, Ti, V, Cr, Mn, Fe, Co, Ni, Cu, Zn, W, Hg, As, Pb, Bi, Se, Th, Rb, U, Sr, Zr, Mo, Ag, Cd, Sn, Sb, Ba, Au, Pt, Pd 3 year warranty | 1
77 | Auto levels white aluminium tripods | Magnification: 20x Standard deviation for 1 km: 2.5 cm Aluminium tripods Aluminium Telescoping Rod Survey Grade: 5 m | 1
78 | Portable Pyrometers | Measuring ranges: 300 - 1999 °C Sensor: Photo diode Measurement area: Ø 10 mm at 500 mm | 1
79 | Table top magnifying lens | Table-mounted magnifying lamp with single glass lens Mechanical double arm minimum 50 cm working radius Minimum 15 cm lens diameter Magnifying power minimum 2.5X Fluorescent white lamp mounted near the lens | 4
80 | Digital Laser Power and Energy Meter | 100pW up to 250W Optical Power range 3 microJ up to 15 J Optical Energy Photodiode sensor input (current): - Accuracy ±0.2% of Full Scale (5microA-5mA) - Display Resolution 1 pA USB 2.0 Interface Numerical and Bar-Graph display Configurable sub-display(s) Simulated Analog Needle for Laser tuning applications Power vs. Time Plotting Beam statistics Direct Optical fiber input 1 GB SD Memory Card Photodiode Power Sensor: - Si Photodiode detector - 9-10 mm aperture - 400-1100 nm WAvelength Range - Power Range 500 nW up to 500 mW - 10 nW or less Resolution - ±0.5% Linearity | 1
81 | Fabry-Perot Benchtop Laser Source | 635 nm, 2.5 mW, FC/PC Single Mode FC/PC Fiber Interface Low Noise stable output Stability: ±0.05dB within 15 min. Setpoint resolution 0.01 mW Adjustable Range: 0-Full Power Modulation Input 0-5 V Modulation Bandwidth: 5 kHz Output Fiber Connector FC/PC Wide 2.1 mm Key Compatible | 1
82 | LED head lamp | Light output: High - 196 lumens; Low - 3 lumens Runtime: High - 5 hours 45 minutes; Low - 75 hours 30 minutes Beam distance: High - 109 m; Low - 17 m 3 X AAA batteries (included) Z Cord in headband for added durability Adjustable elastic straps Easy to operate front power switch Impact and water resistant | 5
83 | 3D camera system | System including: 2 Full HD 11MP cameras with: f/2.8 Fixed Focus, 170º Wide FOV (Including 1080p), 127º Medium FOV (In 1080p) 90º Narrow FOV, 1080p: 1920×1080, 30FPS, 960p: 1280×960, 48FPS + 30FPS, 720p: 1280×720, 60FPS + 30FPS WVGA: 848×480, 120FPS + 60FPS High Performance,1 /2.3” CMOS Image Sensor Light Sensitivity: .84 V/lux-sec, Video Format: H.264 codec, .mp4 File Format Exposure Control: Spot, Center Weighted, White Balance: Auto, Mono, 48 kHz, AAC Compression, Auto Gain Control, Stereo External Microphone Input (3.5mm), 1100mah Rechargeable Lithium-Ion Accessories: 2 Waterproof Housing (197′ / 60m), 2 HD Skeleton Backdoor, 1 Vented Helmet Strap, 1 Head Strap, 4 Curved Surface Adhesive Mounts, 5 Flat Surface Adhesive Mounts, 2 Three-Way Pivot Arm, 1 Suction Cup Mount, 1 J-Hook Buckle, 2 Rechargeable Li-ion Battery, 2 USB Cable 3D System Accessories (For using with 2 cameras): 3D Waterproof Housing, 3D Waterproof Door + Skeleton Door, 3D Sync Cable, 3D Anaglyph Glasses (3 pairs), 3D Helmet Front Mount, 2 Flat + 2 Curved Adhesive Mounts, Assorted Mounting Hardware, Software for 3D movies 2 LCD for cameras with: 2 Waterproof LCD backdoor for HD HERO housing, 2 Waterproof LCD backdoor for HD Wrist housing, 2 Non-waterproof LCD Skeleton backdoor for HD HERO housing, 2 Non-waterproof LCD Skeleton backdoor for HD Wrist housing | 1
84 | hand-held laser range finder | Range: 10 cm - 100 m | 1
85 | Portable LIBS System | Portable LIBS analyzer with ND:Yag laser Spoectrometer range 200-600 nm Optical resolution HR=0.1 nm Detector 2000+ pixels CCD array Pulsed Nd-YAG laser 1.06 microns, 25 mJ Plasma chamber cca 5x5x7 cm Expandible spectrometer chanels Included spectral identification software USB-2 port | 1

## Lot5- Chem. Spectrometry Equ.
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 5: Chemical Spectrometry Equipment
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | FT – IR Spectrophotometer | Detector: High-Performance - Deuterated Triglicine Sulfate (7800-350 cm-1) Beamsplitter: KBr / Ge coated (7800-350 cm-1) Autoalign energy optimization Light Source: MIR source - an ceramic, refractory composite Signal-to-noise: 35000:1 peak to peak in one minute Spectral resolution: 0.4 cm-1 standard resolution Wavelength precision: better than 0.01 cm-1 at 2000 cm-1 Maximum speed: 40 spectra per second at 16 cm-1 resolution Ordinate linearity (ASTM E1421): < 0.1%T deviation from 0.0 %T, measured at 4 cm-1 resolution Electronics: LED control status, PC Connection Frequency Calibration: HeNe laser tube, does not require optimization or calibration Humidity and Vapor Protection: • Sealed and desiccated optical bench with protective BaF2 coating on KBr windows • Rechargeable desiccant cartridges • Humidity indicator Spectrometer Performance Validation: • Integrated validation wheel with NG-11 and NIST traceable polystyrene film standards, serialized • SPV) software ensures spectrometer is performing as expected, day after day (System operation; System Suitability including sampling accessory; Spectral Quality Checks) Software description: • Data collection, Data processing, Data conversion and corrections (Kubelka Munk, KramersKronig, Photoacoustic, ATR correction including settings forcrystal material, reflections, angle and sample refractive index) • Spectral Search : over 9000 spectras • SOP workflows • Spectral peak picking, audit trails history log, multi-format converters for J-camp DX other vendors’ spectral files anduser libraries • Spectral groups statistical analysis: Variance, average, and range • Peak analysis tools • Automatic atmospheric suppression (no standards needed) • QCheck for spectrum-to-spectrum or spectrum-to-multiple spectra QC/QA verification • Quantitative and discriminant analysis package including:Beer-Lambert calibration and prediction, CLS,Discriminant analysis,PLS,PCR ATR accessory: • For analysis of solids, liquids, pastes and gels Installation and training for use. 3 year warranty | 1
2 | Polarimeter | Measuring range of optical rotation: +/- 180° degree Reading: 0.05 degree Monochromatic Light Source: 589 nm (sodium lamp) Included: 1 pc of disc polarimeter 1 piece of sodium lamp 1 piece of 100mm test tube 1 piece of 200mm test tube 4 pcs of test tube protective plate 3 year warranty | 1
3 | Refractometer | Hand held Digital Measurement Range: Brix 0.0 to 50.0% Accuracy: Brix ±0.2% Measurement Temperature: 10°C - 60°C Power supply: batteries 3 year warranty | 1
4 | UV-VIS Spectrophotometer | Optical Design: Dual Beam - Internal Reference Detector Spectral Bandwidth: 1.0nm Light Source (Typical Lifetime): Xenon Flash Lamp Detectors: Dual Silicon Photodiodes Wavelength Range: 190-1100nm Photometric Linear Range: >3.5 A Stray Light: <0.08%T at 220, 340nm (NaI, NaNO2) <1.0%T 198-200nm (KCl) Resolution (Toluene in Hexane): >1.5 PC Control Software description: • System control • Data acquisition • Automatic Data Archiving • Data Management and Custom Reporting • Export data to other applications • Control multiple instruments with one common software platform Desktop data station with monitor and printer: minimium: • Intel Pentium Dual Core E5300 2.6GHz • 2GB DDR2 PC6400 800MHz • integrated graphical and sound card • DVD+/-RW Dual Layer Black • 500GB 7200rpm SATA II • MIDI Tower ATX 450W Black • 22.5“ TFT or LED or LCD monitor • laser printer Installation and training for use. 3 year warranty | 2
5 | XRF Spectrometer | Bench-top instrument Energy dispersive X-ray spectrometer. Air cooling X-ray tube, 5 - 50 kV, 1 - 1000 micro A (ampers) Automatic filter exchanger with at least five filters. Automatic collimator diameter exchanging. Solid state detector with N2 cooling Available for 11Na through 92U quantitative and qualitative element analysis. Concentration range -- low ppm. Noise reduction through automatic changing primary filters. Sample compartment able to handle min. 300m (diameter) x 140 mm (height) samples. Possible vacuum unit. FP model for accuracy measurement, thin film FP method, background correction, matrix correction. Appropriate PC unit with complete software for control and data manipulation. Installation and training for use. 3 year warranty | 1

## Lot6- Microscopes and Acc.
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 6: Microscopes and Accessories
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | Stereo microscope with integrated camera | Compact optical zoom stereo microscope Sudent-proof all-in-one plug-and play design Zoom approx. 8-35X Wide field of view Free working distance at least 90 mm Non-removable eyepieces with diopter setting and variable interpupillary distance Beam splitter 50&50% between camera and eyepieces Integrated full colour digital camera (minimum 3 MP) with independent storage capacity (integrated SD card slot) Live HD function with remote control LED illumination minimum 6500K, providing incident, oblique and transmitted light Variable light intensity control No thermal radiation on specimen site Dust cover, USB cable, HDMI cable, SD card, Imaging software 3 year warranty | 2
2 | Benchtop stereo microscope | Compact optical stereo zoom microscope Zoom 4x with total visual magnification up to 30x Wide field of view Free working distance at least 90 mm Eyepieces 10x/20 and 15x/13 with diopter setting Transmitted light darkfield LED illumination providing incident, oblique and transmitted light Video adapter 3 year warranty | 14
3 | Benchtop stereo microscope | Compact optical stereo zoom microscope Magnification approx. 10x to 80x Supplementary lense for additional 2x magnification (20...160x) Free working distance at least 30 mm Field of view min. 20 mm Viewing angle ergonomic 35 degrees (approx.) Eyepieces 10x and 16x wide field with diopter setting LED illumination providing incident, oblique and transmitted light Banch-top stand with minimal footprint 400x300 mm Sliding stage min. 150 mm diameter Ring epi-illuminating cold light approx. 20W Video adapter 3 year warranty | 1
4 | Stereo microscope with mounting arm | Compact optical stereo zoom microscope Zoom approx. 4x with total visual magnification up to 30x Wide field of view Eyepieces 10x/20 and 15x/13 with diopter setting LED illumination 5600 K providing incident, oblique and transmitted light Integrated coaxial light Horizontal, vertical and tilt movement and repositioning Working distance up to approx. 150 mm 3 year warranty | 1
5 | Biological trinocular microscope | Biological trinocular compact optical microscope Robust metal construction without plastic parts Quadruple nosepiece Mechanical stage with refocusing mechanism Illuminator base min 30 W with field lens with diaphragm Eyepieces 10/20x Abbe condenser Blue filter Objectives: plan achromat 4x, 10x, 40x, 100x oil 3 year warranty | 1
6 | Binocular compact optical microscope | College-grade binocular optical microscope Built-in LED illumination with long lifetime Eyepieces 10x integrated in the tube Infinity corrected plan objectives 4x,10x,40x,100x oil and 100x dry Precentered and prefocused Abbe condenser with slot for Phase and Darkfield Sliders Rounded stage edges with non-extending rack Surface of microscope coated for disease prevention Self adjusting focus mechanism One handed slide loading. Automatic “shut-off” function 3 year warranty | 10
7 | Polarizing microscope | Polarizing trinocular microscope Adaptable eyepieces Photo-tube adapter with T-mount Built-in Bertrand lens (centerable, focusable, retractable) Rotating table 180o with vernier reading to 0.1o First order red plate Three-position interference filter slider Rugged and stable construction Chemical and scratch-resistant surface coating Minimal objective travel range 30 mm Achromatic condenser with iris diaphragm and graduated rotatable polarizer Flatfield achromatic strain-free objectives 4.7x0.11, 9x0.20, 21x0.40, 40x0.65, 95x1.25 oil Easily changeable eyepieces: 10x18, 10x18 with cross-hairs, 6.3x20 with grid scale Reflected light epi-illuminator variable halogen Transmited light illuminator variable halogen Filter set Dust cover Spare light bulb for each light source 3 year warranty | 1
8 | Inverted optical microscope | Compact desktop microscope Integrated bright field and phase contrast 25-40W Halogen lamp and/or LED illumination Set of infinity-corrected Plan Achromat objectives (4-40 X) Eyepiece 10x/20 with diopter setting Tube inclination 45o Adapter for digital camera with 50%:50% beamsplitter Video adapter 0.5 x LD condenser 0.3 Free working distance minimum 70 mm 3 year warranty | 2
9 | Digital microscope camera | 5 MP full colour CMOS camera Basic resolution 2500+ x 1900+ pixels Sensor size minimum 7 mm diagonal Live frame rate 800x600, Frame rate min. 10 Integration time 10 microseconds up to 2 sec Interfaces: SD card slot, mini USB, AV (S-video), DVI-D (HDMI) Spectral range 400-700 nm IR filter C-mount connection Power conncetion via USB 2.0 Supporting OS: Windows XPx32 Prof., Windows 7 3 year warranty | 1
10 | Protective cabinet for microscope slides | Wooden table-top cabinet Capacity of minimum 3,000 standard glass slides Locking doors Metal trays | 1
11 | Compact microtome | Universal tabletop professional biomedical microtome. Enclosed micrometer feeding system which prevents debris to enter the roller bearings Smooth running sledge lockable in at least five positions Precise 8 degrees XY specimen orientation with defined zero position and click stops every turn Section thickness setting: 0.5 to 60 µm Manual feed Clockwise or counter clockwise operation Automatic feeding up to 30 µm Quick release system for specimen clamps Disposable blade holder for low profile disposable blades Safe blade handling Central blade clamping Section waste tray, tool set, dust cover, finger guard Dry ice tray with adapter and thermal gloves. 3 year warranty | 1
12 | Micromanipulator | Single-hand mechanical micromanipulator for stereo or inverted microscope Applicable and ready to be fixed on the most standard stereo microscopes (such as OLYMPUS, NIKON, ZEISS, LEICA) Mechanical three-axis at least 15mm rough pre-positioning Fine single-axis tuning by hydraulic mechanism 2mm total path, 0.2 micrometer minimum graduation Electrode holder 3 year warranty | 2

## Lot7 -Confocal Microscope
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 7: Confocal Microscope
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | Confocal microscope | Motorized confocal inverted microscope Motorized functions: z-focus (10 nm or less), DIC condensor with minimum 5 positions, flurescent cubes handler with minimum 5 positions Fully software controlled instrument with integrated laser-light safety protection. Manually controlled XY microscope table with minimum 50 mm hole. LSM scanning unit with high-speed/high/accurate double galvanomirror system able to support laser light spectrum 350-1100 nm. LSM with three installed lasers (near UV, blue and green) with acoustooptic filter system for fast (2 microsecond or faster) laser control + 3 confocal detectors + external brightfield&DIC detector. LSM able to follow and control laser light intensity while time-laps experiments. Cable-connected fluorescent light source (no microscope frame heating). Possible combined confocal + DIC microscopy Able for classic bright field and reflected fluorescent microscopy Integrated PC computer with 30" monitor Software for image measurements, 2D-3D-4D model rendering with measurements, working focus span setting, objective parafocal parameter tuning, spot measurement, free-line scanning, horizontal and vertical scanning, scanning in selected time intervals, selection of specific area of the image field, optical zoom control (1x-50x with step of 0.1 or less); Simultaneous 4-channel image acquisition (3 confocal fluorescence plus transparent light); 10-60x apochromatic objectives (corrected for continual light spectrum 400-1100nm); DIC equipment for all objective unites Filters for classical fluorescence -- DAPI, FITC nad TRIC. Ready to be upgraded to multiphoton confocal system. 3 year warranty Installation and training for use | 1

## Lot8 -Motorhome for Field Res.
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 8: Motorhome for Field Research
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | Motorhome vehicle | Motorhome vehicle suitable for 5 persons. Maximum weight under 3,300 kg. Two axes chassis. Motor: Diesel engine, 2000/2500 ccm, 100/130 kW, turbo. Fuel tank: cca 100 l. Overall length 7.0 - 7.3m ABS, EBD, Airbag for driver, Servo steering. Spare wheel and tyre. Lap and diagonal seat belts on all designated passenger seats. Double insulated floor. Central locking mechanism. Central heating and ventilation. All walls with thermo insulation. Windows with mosquito net. Central storage compartment 3 cubic meters. Five beds. Internal kitchen with 150 l gas refrigerator. Burner gas hob with electronic ignition and hob lid. Dishwasher sink INOX with cover. Water heater (gas). Storage with external ventilation for two heating gas bottles. Water tank min. 100 l with level sensor. Waste water tank min. 100 l with level sensor. Cassette toilet (C250) with electric flush. Shower compartment. Central electrical control. Min. 3 220V electrical sockets. Min. 2 12 V electrical sockets. Electric power generator 1.5-2.0 kVA with 220V/50Hz adapter/inverter. Additional battery with 220V inverter min. 80 Ah. Radio+CD+MP3. GPS Navigation Retractable awning minimum 4 m wide x minimum 2.5 m with flexible side sheets. External light. 3 year warranty | 1

## Lot9 -Astronomy Equipment
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 9: Astronomy Equipment
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | Telescope | Optical design for observations in primary or secondary focus Primary mirror: - Diameter - 60 cm - f-number - 3-6 - Optical surface roughness < λ/10 RMS - Material – Zerodur , ULE or Astro-Sitall with aluminium coat and (SiO2/TiO2) overcoat Wynne corrector with 3 lens elements Eyepiece - set (4-6 peaces) Ambient temperature range (Deg.C): -15 to +30 Corrected field of 50 mm in diameter (spot size on field edge < 2 arcsec) Carbon truss tube Motorized primary mirror cover | 1
2 | Telescope mount | Robotized equatorial mount with pier Pointing precision RMS < 5 arcsec Tracking precision RMS <0.2 arcsec in 5 min exposure Loading capacity: 120-250 kg Maximum moving speed >5 arcdeg/s Software for mount control | 1
3 | Observatory dome | Big Observatory Dome Main construction material: steel, aluminium, glass reinforced plastic Colour: white Diameter: 6 meters Dome basis with entrance door, 1-1.5m height Fully motorized (dome azimuth rotation, opening shutter) and software controlled (including motors and software) Weather Tight Shutter Operation | 1
4 | Observatory dome | Small Observatory Dome Main construction material: rolling steel, aluminium, glass reinforced plastic or resin coated polyurethane Diameter: 3 meters Dome basis with entrance (1-1.5m) Motorized dome azimuth rotation (including motors and software) | 1
5 | Mobile telescope | Schmidt or Schmidt-Cassegrain Mobile telescope. Primary Mirror Diameter - 25 cm f-number 8 Equatorial computerized mount Eyepiece - set (4-6 peaces) Finderscope Tripod Software for controlling the mount with database with stars, galaxies, deep sky objects, planets, Sun, Moon, and other celestial body objects. With the "Go to" option | 3
6 | CCD camera | CCD camera Astronomical CCD camera Front illuminated Non-antiblooming gate (NABG) type Pixel size 9 micrometers Chip dimension - diagonal > 30 mm (resolution > 6 Mpix) Sensitivity – quantum efficiency > 60% (in peak) Electronic cooling (25 degrees below ambient temperature) With appropriate colour filter wheel (5 filters places) Possibility to connect on telescope tube Connector for computer with data transfer possibility | 2
7 | Photometric filter set | Standard photometric (Bessel) filter set (UBVRcIc) - diameter 50.8 mm | 3
8 | Motorized focus | Motorized focus with software control Loading capacity > 3 kg Ready to be used with main telescope instrument (Item No. 1) | 1
9 | Spectrograph | Astronomical Spectrograph with fiber optic feed Focal Length: 250 mm Wavelength Accuracy: 0.4 nm F-number: 4 Usable Wavelength Range: 200 nm to 25 µm Wavelength Precision: 0.1 nm Gratings: 2400 l/mm, 1200 l/mm, 300 l/mm Micrometer driven variable slit – width from 4 micro- to 3 millimeters, and 2 to 15 mm in height Electronic control of grating carousel | 1
10 | Spectrometer | Handheld spectrometer CCD detector, linear, min. 3000 pixels Detector spectral range 200-1000 nm Sensitivity: max 150 photons/count at 400 nm, max 150 photons/count at 600 nm fiber optic input Signal-to-noise ratio: 300:1 16 bit A/D resolution USB port min. 400 Mbps | 1
11 | Digital planetarium projector | Digital Planetarium projector DLP type of projector Dome diameter 6-8m Resolution (pixel diameter): 1000 Field of View (degrees): 180 Angle of Projection (degrees): ~175 Base Projector Brightness: 2000 Lumens Lamp life: 2500 hours Internal Storage: 1000 GB Lens features: User adjustable focus and centering User Interface: interaction with the projector with a hand-held remote control Number of Stars: 100,000 Software: The planetarium software, Image and video display. Multimedia Support: support for external audio and video connection, full 5.1 channel surround sound | 1

## Lot10 -Glassware and Lab. Sup.
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 10: Glassware and Laboratory Supplies
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | Ag/AgCl electrodes | Reference: Double, Ag/AgCl Connector: 4 mm Banana Connector Electrolyte: 3.5 M Kcl Max pressure: 0.1 bar Temperature range: -5 to 100°C Body material: Glass | 5
2 | Apparatus for electrophoresis | Gel dimensions: 10x8cm 2 gel caster Cooling system Safety lid with leads ceramic plates for gel sandwiches 10 sets of plates (glass+ceramic), with 2 x 2 sets of spacers(0.75 and 1 mm), with 8 combs( 10 samples or more), 0.75 and 1mm thickness; | 1
3 | Apparatus for electrophoresis | Gel dimensions: app 20x15cm and 20x8cm Capable of simultaneously running 2 or 4 gels Cooling system Safety lid with leads Gel casting stands with levelling system, for 2 or more gels Lower chamber for 2L or more of buffer 3 sets of 0.75, 1 and 1.5mm spacers 3 sets of combs (0.75, 1, 1.5) with 10, 15 and 20 wells | 1
4 | Automatic titrator | Ranges: • -2000.0 to 2000.0 mV • -2 to 20 pH • 0 to 100 °C Resolutions: • 0.1mV • 0.1/0.01/0.001pH • 0.1 °C For 5, 10 and 25 mL burette sizes | 1
5 | Digital data acquisition system | Sensor to computer interface Ram min: 512 kb Analog and digital I/O Sensors for voltage, current, thermocouples Software Ethernet: 10 base-t | 3
6 | Hemispherical Heating Mantles | Flask capacity: 2,000 mL Maximum flask diameter approx: 170 mm Power approx: 500 W | 2
7 | Hemispherical Heating Mantles | Maximum flask capacity: 1,000 mL Maximum flask diameter: 130 mm Power approx: 400 W | 2
8 | Hemispherical Heating Mantles | Flask capacity: 500 mL Maximum flask diameter: 100 mm Power approx: 300 W | 2
9 | Hemispherical Heating Mantles | Maximum flask capacity: 100 mL Maximum flask diameter: 60 mm Power approx: 250 W | 2
10 | Iodide ion-selective electrode | Measurement Range: 1.0M to 1X 10-7 M (127000 to 0.11ppm) pH Range: 2 to 13 Temperature Range: 0 to 80°C Connector: BNC | 1
11 | Silver ion-selective electrode | Measurement Range: Ag+: 1.0M to 1X 10-6M (107900 to 0.11ppm) S2- : 1.0M to 1X 10-7M (32100 to 0.003 ppm) pH Range: Ag+ 2 to 8 S2- 12 to 14 Temperature Range: 0 to 80°C Connector: BNC | 1
12 | Nitrate ion-selective electrode | Measurement Range: 1.0M to 1X 10-5 M (62000 to 0.62 ppm) pH Range: 3 to 8 Temperature Range: 0 to 40°C Connector: BNC | 1
13 | Carbon dioxide and carbonate ion-selective electrode | Measurement Range: 1x10-2 M to 1X 10-4 M (440 to 4.4 ppm) pH Range: 4 to 6 Temperature Range: 0 to 40°C Connector: BNC | 1
14 | Calomel electrode | Isopotential Point: pH 7 Temperature Range: 0-60 °C pH Range: 0-14 Connector Type: BNC Calomel Reference Filling Solution (4 M KCl) | 5
15 | Pt - electrode | Electrode for general use including water testing and quality control | 5
16 | Glass electrodes | Temp range: Up to 60°C Electrode connectors: BNC Viscous samples,slurries,suspended solids,sludges Reference cell type: Ag/AgCl pH range: 0-14 | 4
17 | Electrode carrier | For standard electrodes Pivots: 360° Support stand approx: 50 cm | 10
18 | Stirring bar remover | For all stirring bars up to 80 mm in length, PTFE-coated | 5
19 | Magnetic stirring bar set | Length: 2.5 to 80 mm Min 20 pcs of different length PTFE-coated | 3
20 | Magnetic stirring with heating | Max speed: 1500 rpm Electronic speed control with continuous regulation Maximum temperature: 350ºC Indicative diameter of heating plate: 150mm Stirring capacity: 10 l Chemical resistant heating plate | 10
21 | Overhead stirrer | Speed control Max speed: 2000 rpm Stirring capacity: 8 l | 3
22 | Mini Centrifuge | Speed: 12000 rpm Min Capacity: 12 x 1.5 ml tubes | 6
23 | Drying oven | Temperature range: 5 to 300 °C Resolution: +/- 10 °C Interior volume approx: 250 l Stainless steel interior Number of racks: 2 | 3
24 | Laboratory water baths | Temperature range: +5 to 95°C Temperature Uniformity: ±0.5°C Capacity approx: 30 l Bath openings: 130 mm / 4 pcs Sets of rings for openings | 3
25 | Viscometer | Viscosity range: 0.3 to 10,000 cp Software Sample cups RS-232 cable | 1
26 | Distillation apparatus | Unger-type essential oils | 3
27 | Distilling receiver | Vacuum NS 14 | 2
28 | Two neck round bottom flask | 1L NS 29, NS 14 | 5
29 | Two neck round bottom flask | 500 mL NS 29, NS 15 | 5
30 | Round bottom flask | 1 L NS 29 | 5
31 | Round bottom flask | 10 mL NS 14 | 5
32 | Round bottom flask | 100 mL NS 14 | 5
33 | Round bottom flask | 100 mL NS 29 | 15
34 | Flat bottom flask | 1 L NS 29 | 5
35 | Round bottom flask | 25 mL NS 14 | 10
36 | Round bottom flask | 250 mL NS 29 | 15
37 | Flat bottom flask | 250 mL NS 29 | 5
38 | Round bottom flask | 5 mL NS 14 | 5
39 | Round bottom flask | 50 mL NS 14 | 10
40 | Round bottom flask | 50 mL NS 29 | 5
41 | Round bottom flask | 500 mL NS 29 | 15
42 | Flat bottom flask | 500 mL NS 29 | 5
43 | Three neck round bottom flask | 1L 3 x NS 29 | 3
44 | Three neck round bottom flask | 1L NS 29 2 x NS 14 | 3
45 | Three neck round bottom flask | 100 mL NS 29 2 x NS 16 | 3
46 | Three neck round bottom flask | 2L 3 x NS 29 | 2
47 | Three neck round bottom flask | 25 mL 3 x NS 14 | 3
48 | Three neck round bottom flask | 250 mL NS 29 2 x NS 15 | 3
49 | Three neck round bottom flask | 50 mL 3 x NS 14 | 3
50 | Three neck round bottom flask | 500 mL NS 29 2 x NS 15 | 3
51 | Automatic burette, transparent and light | Transparent 10mL | 3
52 | Automatic burette, transparent and light | Transparent 25mL | 3
53 | Automatic burette, transparent and light | Transparent 50mL | 3
54 | Schellbach's burette | 10 mL | 15
55 | Schellbach's burette | 25mL | 15
56 | Schellbach's burette | 50 mL | 15
57 | Screw cap bottle | 100 mL | 10
58 | Screw cap bottle | 250 mL | 10
59 | Screw cap bottle | 500 mL | 10
60 | Beaker | Low form 1 L | 10
61 | Beaker | Tall form 1 L | 10
62 | Plastic beaker | Tall form, 1 L | 10
63 | Beaker | Low form 100 mL | 50
64 | Plastic beaker | Low form 100 mL | 25
65 | Beaker | Low form 2 L | 5
66 | Beaker | Low form 25 mL | 40
67 | Plastic beaker | Low form 25 mL | 20
68 | Beaker | Low form 250 mL | 50
69 | Plastic beaker | Low form 250 mL | 20
70 | Beaker | Tall form 250 mL | 30
71 | Plastic beaker | Tall form 250 mL | 20
72 | Beaker | Low form 50 mL | 40
73 | Plastic beaker | Low form 50 mL | 40
74 | Beaker | Low form 600 mL | 50
75 | Plastic beaker | Low form 500 mL | 20
76 | Dewars flask | Low form Diameter 10 cm | 1
77 | Dewars flask | Low form Diameter 15 cm | 1
78 | Desiccator | Glass Vacuum with stopcock 150mm | 1
79 | Desiccator | Glass Vacuum with stopcock 100mm | 1
80 | Desiccator | Glass Stopcock 200mm | 1
81 | Desiccator | Glass Stopcock 250mm | 1
82 | Graduated centrifugal tubes | 15 mL | 100
83 | Test tubes | 160x16mm 100 pcs | 5
84 | Test tubes | 130x14mm 100 pcs | 5
85 | Flask erlenmeyer | 100 mL | 40
86 | Flask erlenmeyer | 100 mL NS 29 | 40
87 | Flask erlenmeyer | 1000 mL | 10
88 | Flask erlenmeyer | 250 mL | 50
89 | Flask erlenmeyer | 250 mL NS 29 | 20
90 | Flask erlenmeyer | 50 mL NS 14 | 30
91 | Flask erlenmeyer | 500 mL | 25
92 | Flask erlenmeyer | 500 mL NS 29 | 15
93 | Büchner (vacuum) flask | 1 L | 5
94 | Büchner (vacuum) flask | 100 mL | 5
95 | Büchner (vacuum) flask | 2 L | 2
96 | Büchner (vacuum) flask | 500 mL | 5
97 | Rubber septa | NS 14 | 50
98 | Rubber septa | NS 29 | 50
99 | Pasteur pipette bulbs | Capacity: 2 mL Rubber | 50
100 | Hoffman clamps | Pinchcocks 12mm | 10
101 | Hoffman clamps | Pinchcocks 17mm | 10
102 | Hoffman clamps | Pinchcocks 20mm | 10
103 | Hoffman clamps | Pinchcocks 25mm | 10
104 | Needles for plastic syringes | 0.9 mm-100 pcs 0.7 mm-100 pcs | 10
105 | Calcium-chloride tube | Length: 180 mm NB 29/32 | 10
106 | Dropping funnel | 100mL | 3
107 | Dropping funnel | 25mL | 3
108 | Dropping funnel | 50mL | 3
109 | Capillary tubes set | 1/2/3/4/5 µL - Set of 250 pcs 10 µL - Set of 250 pcs 20 µL - Set of 250 pcs | 1
110 | Keck clips | NS 14 | 10
111 | Keck clips | NS 29 | 10
112 | Wire gauzes | Ceramic centre Dimensions: 160 x 160 mm | 10
113 | Kjeldahl flask | 50mL | 15
114 | Kjeldahl flask | 100 mL | 15
115 | Muffles | 16,5 mm | 30
116 | Clamp for burette | Single burette clamp Nickel plated Capacity: 0-40 mm Jaws plastic coated | 15
117 | Universal clamp | Capacity: 0-80mm | 30
118 | Graduated pipettes with piston | Glass 1mL | 10
119 | Graduated pipettes with piston | Glass 2mL | 10
120 | Graduated pipettes with piston | Glass 5mL | 10
121 | Graduated pipettes with piston | Glass 10mL | 10
122 | Knee-shaped connection glass tube | Angle 75 | 5
123 | Knee-shaped connection glass tube | Angle 90 | 5
124 | Allihn condenser | Refluxing Length 15-20 cm 2 x NS 15 | 5
125 | Allihn condenser | Refluxing Length 40-50 cm 2 x NS 30 | 5
126 | Liebig condenser | Length 10-15 cm 2 x NS 14 | 5
127 | Liebig condenser | Length 40-50 cm 2 x NS 29 | 5
128 | Laboratory Jacks | Base approx: 140 x 140 mm Working height: 25 cm | 2
129 | Büchner funnel | Dia 97mm | 5
130 | Büchner funnel | Dia 130mm | 5
131 | Büchner funnel | Dia 156mm | 5
132 | Filter funnel for fast filtration | 75mm | 10
133 | Filter funnel for fast filtration | 55mm | 10
134 | Filter funnel for fast filtration | 100mm | 10
135 | Filter funnel with short tube | 75mm | 10
136 | Filter funnel with short tube | 100mm | 10
137 | Filter funnel with short tube | 125mm | 10
138 | Separating funnel | 1 L | 5
139 | Separating funnel | 100 mL | 5
140 | Separating funnel | 250 mL | 10
141 | Separating funnel | 500 mL | 5
142 | Distilling receiver - bent | NS29 | 5
143 | Distilling receiver | 2x NS14 | 5
144 | Distilling receiver | 2x NS29 | 10
145 | Laboratory coats | L: 20 pcs XL: 40 pcs XXL: 40 pcs PE/Cotton | 1
146 | Laboratory tong | 250mm | 3
147 | Laboratory tong | 220mm | 3
148 | Laboratory tong | 200mm | 3
149 | Measuring cylinder | 10 mL | 40
150 | Measuring cylinder | 100 mL | 20
151 | Measuring cylinder | 25 mL | 20
152 | Measuring cylinder | 250 mL | 5
153 | Measuring cylinder | 50 mL | 30
154 | Measuring cylinder | 500 mL | 5
155 | Microfunnel | Funnel for pycnometer | 25
156 | Microsyringe | 10 μL | 3
157 | Microsyringe | 25 μL | 2
158 | Microsyringe | 50 μL | 2
159 | Volumetric flask | 10 mL | 30
160 | Volumetric flask | 100 mL | 30
161 | Volumetric flask | 1000 mL | 20
162 | Volumetric flask | 2000 mL | 5
163 | Volumetric flask | 25 mL | 50
164 | Volumetric flask | 250 mL | 25
165 | Volumetric flask | 5 mL | 10
166 | Volumetric flask | 50 mL | 25
167 | Volumetric flask | 500 mL | 25
168 | Parafilm | Width: 100 mm Length: 75 m | 5
169 | Parafilm cutter | Cutter/Dispenser Acrylic | 4
170 | Pasteur pipettes set | 1000 pcs | 1
171 | Pycnometer | 100mL | 4
172 | Pycnometer | 10mL | 10
173 | Pycnometer | 25mL | 10
174 | Bulb pipette | 25mL | 10
175 | Graduated pipette | 1 mL | 30
176 | Graduated pipette | 10 mL | 30
177 | Graduated pipette | 2 mL | 30
178 | Graduated pipette | 25 mL | 30
179 | Graduated pipette | 5 mL | 30
180 | Bulb pipette | 50mL | 10
181 | Pipette tilted acc. To Kipp | 5 mL | 3
182 | Pipette tilted acc. To Kipp | 10 mL | 3
183 | Pipette tilted acc. To Kipp | 15 mL | 3
184 | Safety Bunsen Burner | Nickel plated brass Air regulation Pilot flame Lever stop-cock Height: 150 mm For bottle gas, 11 mm | 15
185 | Plastic syringes | 1mL | 5
186 | Plastic syringes | 2mL | 5
187 | Plastic syringes | 5mL | 5
188 | Plastic syringes | 10mL | 5
189 | Porcelain mortar and pestle | Diameter: 150 mm | 10
190 | Safety bulb pipette fillers | Rubber Up to 10 mL | 30
191 | Pipette pumps | 2mL | 5
192 | Pipette pumps | 0,2mL | 5
193 | Pipette pumps | 10mL | 5
194 | Glass atomiser | Glass atomiser with rubber blowball | 1
195 | Reagent bottle | 1 L Transparent and amber | 20
196 | Reagent bottle | 100 mL | 30
197 | Reagent bottle | 250 mL | 30
198 | Reagent bottle | 500 mL | 20
199 | Connection glass tube | 14/23-12/21 | 5
200 | Connection glass tube | 14/23-24/29 | 5
201 | Connection glass tube | 14/23-29/32 | 5
202 | Connection glass tube | 19/26-29/32 | 5
203 | Watch glasses | 120mm | 25
204 | Watch glasses | 40mm | 25
205 | Watch glasses | 50mm | 25
206 | Laboratory hoses | Silicon Diameter: 8/10 mm Length: 50 m | 1
207 | Laboratory hoses | Teflon Diameter: 8/10 mm Length: 20 m | 1
208 | Laboratory silicone grease | 75g | 3
209 | Slotted sieve filter funnel | 50mL | 2
210 | Slotted sieve filter funnel | 75mL | 3
211 | Slotted sieve filter funnel | 125mL | 2
212 | Crystallizing dish | 70mm | 5
213 | Crystallizing dish | 95mm 300mL | 5
214 | Soxhlet Extraction Apparatus with Allihin Condenser | Extractor volume 150 mL | 3
215 | Soxhlet Extraction Apparatus with Allihin Condenser | Extractor volume 250 mL | 2
216 | Spoons and sticks | Plastic or metal Length: 150 mm | 50
217 | Connectors/adapters for rubber hoses | Straight PP For tube inner diameter: 7-9 mm | 20
218 | Glass tubes | Diameter 3 mm Length 1500mm | 5
219 | Glass tubes | Diameter 4mm Length 1500mm | 5
220 | Glass tubes | Diameter 5mm Length 1500mm | 5
221 | Glass tubes | Diameter 6mm Length 1500mm | 5
222 | Glass syringe | 1 mL | 5
223 | Glass syringe | 10 mL | 5
224 | Glass syringe | 2 mL | 5
225 | Glass syringe | 20 mL | 2
226 | Glass syringe | 5 mL | 5
227 | Tubing connectors T and Y shaped | T shaped 10 pcs Y shaped 10 pcs PP For tubes inner diameter: 8 mm | 3
228 | Thermometers | Glass calibrated laboratory thermometer 0 to -50 oC Accuracy 1oC Scale graduation max. 1oC | 5
229 | Thermometers | Glass laboratory thermometer 0 to 300 oC Accuracy 2 oC Scale graduation max. 2 oC | 20
230 | Thermometers | Glass laboratory thermometer NS 14 0 to 360 oC Accuracy 2oC Scale graduation max. 2°C | 10
231 | Thermometers | Glass laboratory thermometer -100 °C through +50 °C Scale graduation max 2 °C | 2
232 | Thermometers | Glass laboratory thermometer -10 °C up to +100 °C Scale graduation at 1 °C | 20
233 | Crucibles | 30x38mm | 15
234 | Crucibles | 40x32mm | 15
235 | Crucibles | 70x56mm | 15
236 | TLC plates | Silica gel 25 x 25 cm Aluminium based | 5
237 | Distilling head | T-shaped 2 x NS 29 1 x NS 14 | 10
238 | Distilling head | T-shaped 2 x NS 14 | 10
239 | Laboratory stands tripod | Inner diameter: 10 cm Height: 18 cm | 10
240 | Triangles for crucibles | Tinned iron wire with clay tubes Length: 60 mm | 15
241 | Tygon laboratory hoses | Diameter: 8/10 mm Length: 10 m | 1
242 | Weighing vessel | 25x40mm | 10
243 | Weighing vessel | 30x50mm | 10
244 | Weighing vessel | 30x60mmm | 10
245 | Distilling column acc. To Vigreaux | Distilling column acc. To Vigreaux length 10cm, 2xNS15 | 5
246 | Distilling column acc. To Vigreaux | Distilling column acc. To Vigreaux length 30cm, 2xNS29 | 5
247 | Safety eye shields | UV filter | 50
248 | Waistcoat Fluorescent | Size: XL | 50
249 | Face Shield | Adjustable headband Dimensions approx: 200 x 300 mm | 50
250 | Test tube racks | For tubes up to 18mm Autoclavable Capacity: 12 tubes | 30
251 | Micro powder spatula | Stainless steel Double ended Length approx: 180 mm | 20
252 | Calibration weights | Stainless steel Class F1 Weights: 1, 2, 5, 10, 20, 50, 100 and 200 g | 2
253 | Support stand | Base plate Support rod length approx: 75 cm | 20
254 | Sand bath | Stainless steel Three block heater | 1
255 | Analytical weighing paper | Grade B-2 Dimensions: 3" x 3" or 4" x 4" 500 pcs | 5
256 | Alcohol tubes | Glass tubes with a neck, dimensions: 37x23mm Plug-type polythene cap 10 pcs | 20
257 | Alcohol tubes | Glass tubes with a neck, dimensions: 40x11mm Plug-type polythene cap 10 pcs | 10
258 | Alcohol tubes | Glass tubes with a neck, dimensions: 50x13mm Plug-type polythene cap 10 pcs | 10
259 | Hand-operated vacuum pump | Plastic Adjustable to vials | 2
260 | Laboratory timer clock | 24 hour count up/down time Loud ring Battery included | 3
261 | Microtubes cooler | Capacity 32 microtubes (1.5 mL) Temperature below −15 °C for up to 2 hours at room temperature Lid filled with gel | 1
262 | Microcentrifuge tube rack | Polypropylene Autoclavable Two rows Capacity: 10 tubes per row | 5
263 | Microscope slide box | Polystyrene Capacity: 100 slides (25 x 75 mm) | 10
264 | Microscoper slide dispenser | Capacity: 50 slides | 1
265 | Wide mouth bottle with screw cap | Made of plastic Colour: white Volume: 50 ml | 20
266 | Wide mouth bottle with screw cap | Made of plastic Colour: white Volume: 100 ml | 10
267 | Wide mouth bottle with screw cap | Made of plastic Colour: white Volume: 250 ml | 1
268 | Polyethylene transfer pipette | Made of polyethylene Bulb draw approx: 3 mL Length: 150 mm Non-sterile 500 pcs | 1
269 | Polyethylene transfer pipette | Made of polyethylene Bulb draw approx: 5 mL Length: 150 mm Non-sterile 500 pcs | 1
270 | Beaker with lid | Made of polypropylene Lid made of polyethylene Volume: 250 ml 100 pcs | 1
271 | Weighing dishes | Made of polystyrene Colour: white Dimensions: 40 × 40 × 8 mm 100 pcs | 1
272 | Weighing dishes | Made of polystyrene Colour: white Dimensions: 86 × 86 × 25 mm 100 pcs | 1
273 | Weighing dishes | Made of polystyrene Colour: white Dimensions: 135 × 135 × 20 mm 100 pcs | 1
274 | Pour boats | Polystyrene Dimensions approx: 40 mm × 30 mm × 10 mm 250 pcs | 1
275 | Pour boats | Polystyrene Dimensions approx: 90 mm × 130 mm × 25 mm 250 pcs | 1
276 | Microscope cleaning kit | Includes: Micro-fiber cloth Lens cleaning tissues Small dust-brush | 5
277 | Turbo Air Electrical Hand Dryer | Wall mounted unit Drying with cold air – no heating unit Anti-microbial external plastics and seals Water ingress protection HEPA filter with anti-microbial coating Touch-free infra-red activation Hand dry time max. 18 sec. Max. power consumption 1600 W 3 year warranty | 8

## Lot11-Biochemistry Equ.
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 11: Biochemistry Equipment
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | SPR Protein Interaction Array System | Number of interaction spots: 20 Refractive index range:1.33–1.37 refractive index units Dynamic range: 1–40,000 RU (response units) Acquisition rate: continuous or 3 Hz (3 images/sec), Baseline drift <1 RU/min, 15–40°C 3 year warranty | 1
2 | Protein LC system | Flow rate: 0.001 - 10mL/min Pressure applicable for protein separation UV-Vis monitor, with 3 simultaneous wavelengths Conductivity monitoring Compatible with all commonly used solvents high quality binary pump Computer controlled Fraction collector Columns: Anion exchange column, 20mg protein capacity, 5MPa pressure limit Cation exchange column, 20mg protein capacity, 5MPa pressure limit Gel filtration column, Sepharose or Superose based, app 10x300mm, at least 1.5MPa pressure limit C18 reverse phase column, app 4.6 x 250mm, 10MPa pressure limit C5 reverse phase column, app 3x150mm, 8MPa pressure limit 3 year warranty | 1
3 | Patch clamp system | Patch clamp setup, containing: -Phase-contrast or Nomarski microscope, invert, 10x, 40x, 100x objectives, with suitable stage for attachment to an antivibration table and flood management - antivibrational table, breadboard with no holes; max. relative tabletop motion: < 1nm -8bit CCD microscope camera withImmunofluorescence port and calcium fluorescence port -Micromanipulators, with three axis movement, with mechanical and electrical( computer operated) controls -Mechanically controlled perfusion system -Micropipettes with filament for electrophysiology (200pcs) -Pipette puller, horizontal model, with tip polishing functions -Patch clamp amplifier - digitally controlled, with appropriate software -Personal computer, minimum 4GB RAM, 2GHz dual core CPU, 1TB HDD, 22.5" flat screen 3 year warranty | 1
4 | Microplate reader | For 96-well microplates Detection mode: absorption Wavelength range 340-750nm Resolution 0.001OD Read method: endpoint, kinetic Computer controlled Multiple microplate types 3 year warranty | 1
5 | Gel documentation system | Real time image 8-bit CCD camera Microprocessor controlled darkroom Integrated colour LCD Instant saturation detection USB port for data transfer 312nm transilluminator 3 year warranty | 1
6 | Electrode puller for intracellular electrophysiology | Environmental chamber Programmable air pressure Pull solenoid power supply rectifier Electrode tip diameter: 0.1 µm Looping pull cycle Pre-programmed sample programs 3 year warranty | 1
7 | Autoradiography cassette | 18x24cm membrane size Temperature up to -70 °C Non-porous pad | 2
8 | Real-Time PCR System | • The QPCR instrument thermal system must have temperature uniformity across the 96-well block of +/- 0.25°C at annealing and extension. • The QPCR instrument must use a quartz tungsten halogen lamp for excitation to deliver light from 350nm to 750nm wavelength range. • The QPCR instrument must have five (5) narrow band-pass excitation filters and five (5) narrow bandpass emission detection filters which are user-selected at time of QPCR system purchase. • The QPCR instrument must use coaxial fiber optic bundles to simultaneously deliver excitation light and detect fluorescence emission light from the dye/sample. • The QPCR instrument must use scanning fiber optics to ensure that all 96-wells receive the same excitation light for the same duration and are the same distance from the detector. • The QPCR instrument must use a photomultiplier tube (PMT) for detection of emitted fluorescence signal to ensure a superior signal-to-noise ratio which leads to increased sensitivity for detection of low copy number targets, ability to discriminate 2-fold differences in target concentration, and detection across ten (10) orders of magnitude of the dynamic range in target concentration. • The QPCR system must support custom filter control so an excitation filter can be matched to a longer or shorter wavelength emission filter. • The QPCR system must support real-time quantitative PCR, comparative quantification, dissociation curve analysis, and quantitative and qualitative plate read experiments, concentration, and detection across ten (10) orders of magnitude of the dynamic range in target concentration. • Desktop/Laptop computer with application-specific software. • The vendor must provide comprehensive assay design and development guidelines for real-time quantitative PCR and SNP genotyping assays. • 3 year warranty | 1

## Lot12 -Env. and Fielad Res. Equ
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 12: Environmental and Field Research Equipment
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | Multiparameter water quality meter | Mobile handheld field instrument Multiparameter integrated probe up to 13 sensors: - Water level - pH: -2.00 to 14.00 pH (accuracy: +/- 0.01 pH) - mV: -2000 to +2000 mV (accuracy: +/- 1 mV) - Temperature: -5.0 to +100.0 °C (accuracy: 0.1 oC) - Conductivity: 0.0 μS/cm to 2000 mS/cm - Salinity: 0.0 to 70.0 - Dissolved oxygen: 0 – 30 mg/L - Oxygen saturation: up to 300% - Nitrate: up to app. 50 g/L, ±1 mg/L - Ammonia: up to app. 15 g/L, ±1 mg/L - Chloride: up to app. 30 g/L, ±1mg/L - Fluoride: up to app. 20 g/L, ±1mg/L - Sodium: up to app. 20 g/L, ±1mg/L - Calcium: up to app. 30 g/L, ±1mg/L - Potassium: up to app. 35 g/L, ±1mg/L - Chlorophyll A: up to app. 500 μg/L - Cyanobacteria: up to app. 2 mill. cells per mL - Resistivity: 0.00 to 2000 Mohm Temperature compensation Standard 2 m cable Watertight housing Rechargable batteries with external charger Alphanumerical 3 line LCD display RS232 external interface RS485 sensor interface Data logger: serial flash-memory with 4MB, max 32 channels, automatic/manual storage | 2
2 | Integrated multi-probe water quality system | Small handheld instrument Simultaneously measurement of dissolved oxygen (0-50 mg/L), pH (0-14, accuracy 0.2 pH), conductivity (0-100 mS/cm), temperature (-5 to 45oC) and ORP (-999 to +999 mV) Impact resistant case Data storage min. 40,000 data sets, time and date Rechargable battery pack with fast charger RS232 cable for external connection to PC 10 m cable pH/ORP Kit Flow Cell | 3
3 | Conductometer | Range: 0.0 to 199.9 µS/cm; 0 to 1999 µS/cm; 0.00 to 19.99 mS/cm; 0.0 to 199.9 mS/cm Resolution: 0.1 µS/cm; 1 µS/cm; 0.01 mS/cm; 0.1 mS/cm Accuracy: ±1% Calibration: manual, 1 point Temperature Compensation Probe: HI 76302W, ATC, with 1 m cable Battery Type / Life: 1 x 9V / approx. 100 hours of continuous use | 3
4 | Laboratory pH meter | Range pH : 0.00 to 14.00 Range mV : ±1999 Resolution pH : 0.01 Resolution mV : 1 Accuracy (at 20°C) pH: ±0.01 Accuracy (at 20°C) mV: ±0.1 pH Calibration: Manual 2 points Temp Compensation Electrode 50-HI1332B: combination double-junction refillable pH electrode BNC with 1 m cable | 6
5 | Precision laboratory balance | Readability: 0.01 g Repeatability: 0.02g Calibration mass: 2000 g | 5
6 | Mobile laboratory for water analysis set | Set includes reagents, apparatus and accessories for determination of: Parameter/Range/No. of Tests Acidity (as CaCO3)/10-4,000mg/L/100 Alkalinity (as CaCO3)/10-4,000mg/L/100 Ammonia (as NH3-N)/0-0.50mg/L/100 Carbon Dioxide/10-1,000mg/L/100 Chlorine,Free/0-2.00mg/L/100 Chlorine,Total/0-2.00mg/L/100 Hardness (as CaCO3)/10-4,000mg/L/100 Iron,Total/0-3.00mg/L/100 Manganese,HR/0-20.0mg/L/100 Nitrate,LR(as N)/0-0.50mg/L/100 Nitrite (as N)/0-0.350mg/L/100 Phosphate/0-2.50mg/L/100 Sulphide/0-0.70mg/L/100 Tannin-Lignin/0-9.0 mg/L/100 Temperature/-10 to 110°C/ Zinc/0-3.00mg/L/100 Caring box | 1
7 | Mobile pH meter | Measures: pH, mV, °C pH Range: 0.00 to 14.00 pH pH Accuracy: +/-0.01 pH pH Calibration: 2 points mV Range: -200 to +200 mV, -2000 to +2000 mV mV Accuracy: +/-0.2 mV, +/-2 mV Temperature Range: -10.0 to 100.0 °C Temperature Accuracy: ±0.5 °C Temperature Calibration | 3
8 | UV TLC lamp | UV source: 365nm, 254nm or split tube 254/365nm | 1
9 | Vortex | Rotation speed: 0 to 3000 rpm Speed setting: analog Speed Control: electronic Operation mode: motion detector or continuous Shaking motion: orbital Orbital diameter: 4.5 mm | 2
10 | Chromatography chambers | Big: Glass For 200 x 200 mm TLC plates Small: Glass or plastic For plates 25 x 70 mm TLC plates | 3
11 | Fused quartz cuvettes | Optical pathlenght: 10 mm Height: 45 mm Volume: 3.5 mL | 4
12 | Fused quartz cuvettes | Optical pathlenght: 10 mm Height: 45 mm Volume: 1.4 mL | 4
13 | PCR machine | 48 x 0.5 ml PCR tubes Temp control range: 4-99 °C Block homogeneity: 0.5 °C Programmable On-device programs:100 Max no. of cycles: 99 | 3
14 | Electrophoresis power supply | Voltage: 6-600 V Current: 1-400 mA Power: 1-100W 3 protocols in memory 500h continuous operation | 2
15 | Minifuge with cooling | Temperature mode: 0-40 °C 30x1.5mL capacity Max speed: 15.000 rpm Short impulse mode(quickspin) | 1
16 | DNA Electrophoresis Lab Set | Lab set includes: Power supply 4 pcs Mini-Sub cell GT systems 4 pcs 2–20 μl adjustable-volume micropipets | 1
17 | Thermostatic circulator | Controls coolant temperature from -10°C to +90°C ± 0.1°C; 3L reservoir; Flow: 10L/min; Noise level: 65dB | 1
18 | Transiluminator | 3 UV wavelength: 254, 302, 365nm UV blocking cover Lamp power: 10watt; | 1
19 | UV lamp | Portable Lamp length app: 90 cm Power app; 15 W 10.000h lamp life Stainless steel casing and mounting | 2
20 | Peristaltic pump | Min flow: 0.5-1 ml/h Max flow: 500 ml/h Automatic pressure plate for convenient tubing changes Reverse flow function | 2
21 | Multichannel pippete | 10-100μL; 0.1μL volume increment; 12 channels, autoclavable | 2
22 | Dewar flask | Volume: 1L Insulated lid Wide mouth Nitrogen loss without lid: 0.1L/hour | 2
23 | Incubator / Orbital shaker | Temp range: 5°C above ambient to 70°C Speed: 15 to 400 rpm Load: 15 kg Clamps for 500mL flasks Shaker platform | 1
24 | Ultrafiltration cell | Process volume 50mL Minimal volume: 2.5mL Effective membrane area 10 cm2 Spare set of O-rings, tubing, and membrane holders | 1
25 | Soil Sampler Kit, Liner (Core) for Medium to Soft Soils | Bailer boring equipment Hand operated Auger set for heterogeneous soils For soil depth of 7 m Unpainted augers for environmental research | 1
26 | Microbial air sampler | Flow rate: 30 lpm Built-in timer Batteries and charger included For 90mm petri dishes Portable | 1
27 | Aquatic temperature datalogger | Waterproof Working and logging range: -10°C to 70°C Resolution: 0.01°C Logging interval: 1 sec - 10 days No. of readings: 30,000 Dimensions approx: 50 x 50 x 50 mm | 2
28 | Bat detector | Time Expansion and Heterodyne detection methods | 2
29 | Temperature and humidity datalogger set | Temperature range: -20°C to 80°C Temperature sensor: 10K NTC Thermistor Temperature resolution: 0.01°C Relative humidity range: 0-100% RH Relative humidity resolution: better than 0.5% RH Memory: 64k No. of readings: 20,000 Logging interval: 1 sec to 10 days | 2
30 | Utermöhl plankton chamber set | Set includes at least: 1 Utermöhl cylinder 10 ml 1 Utermöhl cylinder 50 ml 1 Utermöhl cylinder 100 ml 3 cover plates 50 base plates | 2
31 | Berlese apparatus/funnel | Metal Diameter approx: 20 cm | 2
32 | Sterility indicator strips | Bacillus stearothermophilus spores impregnated on paper 25 pcs package | 4
33 | Entomological pooter/aspirator | Plastic | 30
34 | Compact binoculars | Optical parts 10 x 25 or 10 x 42 | 11
35 | Counting chamber | Neubauer design | 2
36 | Digital night vision binoculars | Optics 5x42 or 10x42 Viewing distance 600 m | 2
37 | Portable dissolved oxygen meter | Range (mg/l Oxygen): 0.00-45.00 mg/L Resolution (mg/l Oxygen): 0.01 mg/L % Saturation Oxygen Range: 0.0-300.0% % Saturation Oxygen Resolution: 0.1% Calibration: automatic in saturated air | 1
38 | Dynamometer scales | Capacity: 300 g Resolution: 2 g | 2
39 | Dynamometer scales | Capacity: 600 g Resolution: 5 or 10 g | 1
40 | Dynamometer scales | Capacity: 2500 g Resolution: 20 or 50 g | 1
41 | Standard entomological net | Diameter: 30 cm | 5
42 | Sweeping entomological net | Diameter: 30 cm | 5
43 | Flexible plastic endoscope | Flexible optical cable, diameter 10mm and length 1 m Lamp: integrated solid state LED Image resolution approx 5000 pixels | 2
44 | Kick-net aquatic sampler | Stainless steel ring/frame Ring/frame opening approx: 25 x 25 cm Net mesh size: 500 μm | 2
45 | Kick-net aquatic sampler | Stainless steel ring/frame Ring/frame opening approx: 25 x 25 cm Net mesh size: 800-900 μm | 1
46 | Small mammal trap | Humane trap Made of plastic | 50
47 | Small mammal trap | Longworth design Made of aluminium or stainless steel | 20
48 | Malaise trap | Standard, black and white colour | 2
49 | Entomological marking paints | Standard 5-colour set Paint: quick drying, non-toxic, water-based | 2
50 | Entomological marking pens | Standard 5-colour set Pens: quick drying, non-toxic, water-based | 5
51 | Plankton counting chamber | Kolkwitz design Volume: 0.5ml | 3
52 | Plankton counting chamber | Kolkwitz design Volume: 1 ml | 3
53 | Plankton net | Mesh size: 18 μm | 2
54 | Plankton net | Mesh size: 25 μm | 2
55 | Portable digital balance | Alkaline batteries and AC adapter included Display: LCD Range: 0-2000g Accuracy: 1.0 g | 3
56 | Portable digital balance | Alkaline batteries and AC adapter included Display: LCD Units: grams Range: 0-200g Accuracy: 0.1 g | 4
57 | Portable luxmeter | Ranges and Resolution: 0.001 -1.999 Klux ± 0.001 Klux; 0.01 - 19.99 Klux ± 0.01 Klux; 0.1 - 199.9 Klux ± 0.1 Klux | 2
58 | Water sampler | Ruttner design Capacity: 1L | 2
59 | Water sampler | Ruttner design Capacity: 3L | 2
60 | Sample container with cap | Plastic Volume: 10 mL 1000 pcs | 1
61 | Sample container with cap | Plastic Volume: 30 mL 1000 pcs | 1
62 | Secchi disc | Standard: EU Norm EN 27027 3/94 | 2
63 | Shaking water bath | Temperature range: 20 to 100°C Shaking frequency: up to 200 cycles per minute, adjustable Inner volume: up to 20 L | 1
64 | Snake catcher | Telescopic handle, made of aluminium | 2
65 | Vibration damping mount | Isolates from vibrations down to 9 Hz Resistant to chemicals Support: 15 kg Dimensions 30 x 30 cm, thickness 10 cm | 1
66 | Soil pH meter | Handheld pH range: 0 – 14 pH resolution: 0.01 Temperature range: 0-90°C | 2
67 | Soil moisture meter | Handheld Relative moisture range: 0 – 100 % Resolution: 1% | 2
68 | Hand tally counter | 4-digit counter | 10
69 | Portable pH/ECTDS/Temp tester | Waterproof Temperature probe made from stainless steel Range; resolution: TDS: 0.00-10.00 ppt; 0.01 pH: 0.00-14.00 pH; 0.01 Conductivity: 0.00-20.00 mS/cm; 0.01 | 2
70 | Binoculars with digital camera | Optical parts: 10 x 25 or 10 x 42 Digital part: 4.0MP 4x Digital Zoom 6 MB RAM Internal Memory SD Card Slot | 2
71 | Photosynthesis measuring system | Handheld Display: LCD Data Storage: 4 MB Internal RAM Data output: USB CO2 Measuring Range: 0 to 2000 ppm, resolution: 0.1 ppm H2O Measuring Range: 0 to 100%, resolution: 0.1% PAR Measuring Range: 0 - 2500 µmol m-2s-1 Temperature Range: - 10 - 50 °C | 1
72 | Standard geological compass | Three inclinometer systems: card inclinometer (±5° accuracy with 5° resolution), hinge inclinometer (±5° accuracy with 5° resolution), and graduated dial inclinometer (±1° accuracy with 1° resolution). Body waterproof, lightweight and rugged. Adjustable declination. Magnified read out plus map magnifier. Exchangeable card scales: USGS 1:24,000 feet, mile, UTM, and meter scales. Accuracy: ±1°. Resolution: 1°. Dimensions: 4.1" x 2.5" x 1". | 8
73 | Brunton-type Geological Compass | Aluminium housing. Induction damped needle with sapphire jewel bearing. O-rings waterproof construction. Adjustable needle locking mechanism with two positions. Hinge inclinometer with 2° increments for 1° readable dip measurements. “Buck Horn” style sights for increased accuracy when sighting azimuth or vertical angles. Horizontal angles: 0-360° (Azimuth) or 0-90-0° (Quadrant), 1° graduations. Vertical angles: ±90°, vernier reading to 10 minutes, ±100% grade, 5% divisions. Declination adjustment. Leather case. | 3
74 | Tandem compass with clinometer | Liquid damped precision compass and clinometer in one compact, rugged aluminium housing protected against impact, corrosion, and water. Clinometer scale is graduated in percent and degree (0-90°, 0-150%). Compass scale is in azimuth (0-360°) with the option of declination adjustment (20° East and West). The compass graduated in 1/2° increments. The clinometer graduated in 1°/1% increments. Both the compass and clinometer are individually calibrated. Threaded tripod socket. Carrying case and lanyard. | 1
75 | Water sampler field kit | Kit should include: clear polycarbonate 1.0 litre water sampler, 200g messenger, 20 cm limnological Secchi disc; armored thermometer (liquid spirit filled, -5°C to +40°C in 0.5° increments); plankton net (20 cm diameter x 60 cm length x 153 µm); three 20.0m nylon cords, line reels, plastic carrying case, and instructions. | 1
76 | Handheld Flowmeter | 1-2m expandable flow probe. Located on the bottom tip is a protected 2“ turbine sensor which uses positive displacement technique. The sensor rotates freely on a bearing shaft with no mechanical interconnections. Magnetic material in the propeller passes a pickup coil in the housing producing electrical impulses. Readout display on top of the handle, feet/meters per second readings. Instantaneous velocity and true average velocity measurement, minimum/maximum velocity, time-of-day, timer, totalizer for batch flows, and RPM for accurate low velocity measurement. Record memory (at least 20 measurements). Gun-type carrying case. Specifications – Range: 0.3 to 15 FPS (0.1 to 4.5 MPS). Accuracy: average velocity ±0.1 FPS, instantaneous velocity ±0.5 FPS. Averaging: true digital running average. Materials: PVC handle and propeller housing, anodized aluminium shaft, brass bearing. Power: battery. | 1
77 | Mineral identification and hardness kit | Test for mineral hardness under Mohs’ hardness scale ranging from 1 to 10. Kit should include four double-ended picks comprising eight points made of various metals and alloys with Mohs’ hardness values numbering 2, 3, 4, 5, 6, 7, 8, and 9. Points large enough to be resharpened. Kit should also include a grinding stone for sharpening. Glass plate with a hardness of 5.5 which, when scratched, quickly determines a hardness range. Streak plate for testing mineral’s streak colour. Magnet for testing magnetic rocks such as magnetite and pyrrhotite. All tools packaged in compact, case. | 1
78 | Geological hammer | Hammer made of a single piece of solid steel. One side with a flat square surface, other in form of a pick with a sharp point end. Polished handle covered with a non-slipping plastic material, rubber or leather. Length: 30-35 cm. Mass: 700-1400 g. | 10
79 | Geological rock chisel set | Set of three chisels: one 200-250 mm long with a point tip, one 250-320mm long with a narrow chisel end point, and one 200-240 mm long with a wider chisel cut point. Each chisel made of a single piece of hard steel. Vinyl or rubber protected handles. | 4
80 | Handheld 7-gas emission analyzer | Integrated built-in electrochemical sensors and CO2 IR bench Integrated condensate trap with PTFE filter Color display USB connection Integrated operating software SD card for data storage Simultaneous gas analyze: - CO2 -- 0-20% (accuracy 5% of meassured value) - O2 -- 0-21 % (accuracy 0.2 vol.%) - NO -- 0-1000 ppm (accuracy 5% of reading) - NO2 -- 0-200 ppm (accuracy 5% of reading) - SO2 -- 0-2000 ppm (accuracy 5% of reading) - H2S -- 0-200 ppm (accuracy 5% of reading) - CO -- 0-4000 ppm (accuracy 5% of reading) Battery operated by Li-Ion battery Two Spare Batteries Transport case | 1
81 | Ground penetrating radar | Ground Penetrating Radar: Data acquisition system: 50.000 samples/sec Single channel Transmitting range: 80 KHz Internal storage Portable Weight approx: 4 kg Caring bag Battery: 10 V Antenna: Frequency: 400 MHz Antenna control cable Rugged four wheel cart: Wheel diameter approx: 15 '' Compatible with radar Internal integrated survey wheel encoder Data processing software whit mapping module 3 year warranty | 1
82 | Field Spectrometer | Portable field instrument with microprocessor and data logger Modular construction ready for further upgrade Light collection lens, cable and accessory Integrated compact spectrometer 200-1050 nm Integrated Java-based software ready to operate on Windows and Linux OS and with full control of spectrometer operations and data storage. SD card slot Battery module Waterproof 3 year warranty | 1
83 | Colorimetry System | Concave grating spectrometer for UV-VIS spectrocolorimetry with integrated filters and interface cable. Reflectance fixture with integrated light source and fiber optic cable Reflectance standard to set white reflectance, 50 mm diameter. 45o angled tip for reflectance probes Reflectance proble holder for 45o angle measurements Reflectance flexible fiber probe for UV+VIS (7 around 1) Real-time opearting software for Windows 9x/XP/VISTA/7 3 year warranty | 1

## Lot13 - Basic Chem and Bio Equ.
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 13: Basic Chemistry & Biology Equipment
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | Analytical balance | Capacity approx: 200g Resolution: 0.1 mg Repeatability: 0.0001 g Pan size approx: 8 cm diameter | 4
2 | Centrifuge with cooling | Capacity: 4x50 mL Max. Speed: 20000 rpm Timer: 0 to 99 h Temperature range: -5 to +40°C, ±1°C Acceleration/Deceleration adjustments Set of adapters for 5, 15, 50 and 85mL tubes | 1
3 | Volumetric Karl Fischer titration | Karl Fischer apparatus For determination of water content according to Karl Fischer Automatic titrator Titration unit Titration stand with integrated stirrer and pump Titration vessel Platinum electrode | 1
4 | Muffle Furnace | Capacity approx: 15 l Temp range: ambient to 1200°C Power approx: 4000 W Door to be used as shelf | 1
5 | Laboratory incubator | Temperature Range: 5°C to +70°C Uniformity: ±0.5°C Capacity: 50 l | 2
6 | Ultrasonic cleaner | Tank capacity approx: 10 L Frequency: 40 KHz Heating up to 80 °C Timer up to 15 min | 2
7 | Vacuum pump | Max vacuum: 29.7" Hg (2 torr) Max pressure: No pressure output Free-air capacity: 2.5 CFM (70.0 LPM) Duty cycle: continuous Max temperature: 40°C Head configuration: Dual | 3
8 | Vacuum evaporator | Bath temperature: 20-180oC Bath Capacity: 3 litres Temp resolution: +/- 1oC Display: Digital Rotation speed range: 20-280 rpm Configuration: Dry Ice Reflux Condenser Vacuum connection: Advanced vacuum controller Glass type: Normal | 1
9 | Melting Point Apparatuses | Temp range: ambient to 300°C Resolution: thermometer dependent Reproducibility: ±1°C | 1
10 | First Aid Box | For chemistry and physics areas | 3
11 | Full face mask | Filters Exhalation valve Size: L | 2
12 | Chemical resistant overall | Size: XL Resistant to acid and alkalies Multi layered PE Hood and zip Welded feet | 2
13 | Safety cans | Capacity: 10 L Flame arrester Lead fuse Screw cap Stainless steel | 1
14 | CO2 incubator | Volume: 100L Automatic sterilization cycle Temp. fluctuation at 37 C: <0.3 C Sterilizable humidity sensor Fault diagnostics, over temperature protection Digital electronic CO2 control with NDIR | 1
15 | Freezer unit | Temp range: -40°C to -80°C Capacity: 500L Upright configuration Alarm system Remote alarm relay Digital temperature control | 1
16 | Antivibrational table | Maximum dynamic deflection coefficient: <0.01 Maximum relative tabletop motion: < 1nm Breadboard with no holes | 1
17 | Laminar flow chamber | Compliance with EN-12469 standards HEPA H14 EN 1822 filters Load capacity 50kg Working height - adjustable 750-950mm Dimensions approx: 2200x1300x900 mm | 3
18 | Liquid nitrogen tank | Liquid nitrogen capacity: 40 l Static evaporation rate: 0.4L/day Normal working duration approx: 75 days | 1
19 | CO2 gas tank | Volume: 50 L | 4
20 | N2 gas tank | Volume: 50 L | 3
21 | Thermoshaker | Interchangeable block modules for micro plates and micro tubes up to 2.0 ml Temp. range: RT+5 C - 100 C Variable shaking speed: 250-1400 rpm Heat up time app 15min Micro plates block Micro tubes block | 3
22 | Micropipette controller | For micropipettes and small volume pipettes up to 1 mL Autoclavable | 5
23 | Centrifuge tubes with cap | Polypropylene Volume 15 mL 500 pcs | 1
24 | Centrifuge tubes with cap | Polypropylene Volume 50 mL 25 pcs | 1
25 | Centrifuge tubes with cap | Polypropylene Self-standing Volume 50 mL 500 pcs | 1
26 | Colony counter | For water-resistant markings Free-standing magnifier (up to 2x) included Removable grid included | 1
27 | Crayfish trap | Vinyl coated steel Frabill design | 10
28 | Cryoclaw manipulator for vials | Stainless steel with rubber cups | 1
29 | Cryogenic storage rack | Stainless steel Vertical, 4 shelves For boxes of approx 3"×3"×2" | 1
30 | Cryogenic vials | Polypropylene Sterile Volume: 2 mL For temperatures to −196 °C 500 pcs | 1
31 | Dissecting kit | Includes: surgical scissors straight surgical scissors curved tissue forceps thumb forceps handles pointed needle arrow needle chain hook Stainless steel | 10
32 | Dissecting metal needle with holder | Pointed Straight Stainless steel | 5
33 | Dissecting metal needle with holder | Pointed Curved Stainless steel | 5
34 | Electrophoresis plate rack | Polypropylene 10 places | 2
35 | Entomological pins | Black enamelled steel Size 0 100 pcs | 10
36 | Entomological pins | Black enamelled steel or stainless steel Size 00 100 pcs | 10
37 | Entomological pins | Black enamelled steel or stainless steel Size 000 100 pcs | 10
38 | Entomological pins | Black enamelled steel or stainless steel Size 1 100 pcs | 10
39 | Entomological pins | Black enamelled steel or stainless steel Size 2 100 pcs | 10
40 | Entomological pins | Black enamelled steel or stainless steel Size 3 100 pcs | 10
41 | Entomological pins | Black enamelled steel or stainless steel Size 4 100 pcs | 10
42 | Entomological pins | Black enamelled steel or stainless steel Size 5 100 pcs | 10
43 | Entomological pins | Black enamelled steel or stainless steel Size 6 100 pcs | 10
44 | Entomological pins | Black enamelled steel or stainless steel Size 7 100 pcs | 10
45 | Eppendorf-like automatic pipette | Adjustable-volume: 0.1-2.5 μL Box for tips included | 5
46 | Eppendorf-like automatic pipette | Adjustable volume: 0.5-10 μL Box for tips included | 5
47 | Eppendorf-like automatic pipette | Adjustable volume: 10-100 μL Box for tips included | 15
48 | Eppendorf-like automatic pipette | Adjustable volume: 100-1000 μL Box for tips included | 15
49 | PCR cooler | Eppendorf-like design | 2
50 | Extractor for ethidium bromide decontamination | Flow: 1 lpm Decontamination capacity: 99% of EtBr 5 pcs | 2
51 | Gel cutter | Acrylic Length: 20 cm | 1
52 | Gel scoop | Acrylic Dimensions: 20 × 20 cm UV-transparent | 1
53 | Dropping bottle with graduated dropper | Glass Capacity: 30 mL | 2
54 | Dropping bottle with graduated dropper | Glass Capacity: 125 mL | 1
55 | Glass spreader for inoculating petri dishes | 90° bend 5 pcs | 1
56 | Glove bag | Polyethylene bag with two-hand openings Closure: zipper-lock Volume: 50 L | 1
57 | Graduated laboratory bottles with cap | Made of clear borosilicate glass Capacity: 50 mL 10 pcs | 2
58 | Graduated laboratory bottles with cap | Made of clear borosilicate glass Capacity: 100 mL 10 pcs | 2
59 | Graduated laboratory bottles with cap | Made of clear borosilicate glass Capacity: 250 mL 10 pcs | 2
60 | Graduated laboratory bottles with cap | Made of clear borosilicate glass Capacity: 500 mL 10 pcs | 1
61 | Graduated laboratory bottles with cap | Made of clear borosilicate glass Capacity: 1000 mL 10 pcs | 1
62 | Inoculating loop for microbiology | Wire made of 19 gauge platinum with 5% rhodium Length: 70 mm Capacity: 1 μl | 5
63 | Inoculating loop for microbiology | Wire made of 19 gauge platinum with 5% rhodium Length: 70 mm Capacity: 10 μl | 2
64 | Inoculating loops for microbiology | Made from polystyrene Sterilized Capacity: 10μl 1000 pcs | 1
65 | Mobile laboratory glassware washer | Movable spring clips can accommodate flasks from 50 to 2000 mL Small spindles can hold flasks from 1 to 25mL Clips for narrow-mouth glassware included Vacuum drying system included Programmable LCD display Dimensions approx: 60 × 70 × 90 cm | 1
66 | Microcentrifuge filters | Capacity: 500 μL Duramore PVDF membrane, pore size (MWCO) 0.22 μm 25 pcs | 3
67 | Microcentrifuge filters | Capacity: 500 μL Duramore PVDF membrane, pore size (MWCO) 0.45 μm 25 pcs | 1
68 | Microcentrifuge filters | Capacity: 500 μL Duramore PVDF membrane, pore size (MWCO) 0.65 μm 25 pcs | 1
69 | Microcentrifuge tubes with lid | Volume: 0.5 mL 1000 pcs | 2
70 | Microcentrifuge tubes with lid | Volume: 1.5 mL 1000 pcs | 2
71 | Microplate shaker | Variable speed control Speed approx: 1500 rpm | 1
72 | Vacuum filtration set | Glass For 47 mm filters Included parts: funnel top, funnel stem/membrane support, aluminium clamp, stopper | 1
73 | Syringe filter unit | Fluoropore PTFE (F) filter membranes in plastic housings Pore size: 0.22 μm Sterilised 100 pcs | 1
74 | Nitrile gloves | size S, 100 pcs | 5
75 | Nitrile gloves | size M, 100 pcs | 10
76 | PCR plate | Polypropylene Colourless Capacity: 48 well 25 plates/pack | 2
77 | PCR sealing film | For 48 well 100 pcs | 1
78 | PCR tubes | Volume: 0.5 mL 1000 pcs | 2
79 | Pestle | Polypropylene For 0.5 mL tubes 100 pcs | 2
80 | Pestle | Polypropylene For 1.5 mL tubes 100 pcs | 2
81 | Pipette brush | Made of natural bristle | 5
82 | Pipette canister | Stainless steel Square cross-section For 0.1-0.5 mL | 1
83 | Pipette canister | Made of stainless steel Square cross-section For 1-10 mL | 5
84 | Pipette canister | Made of stainless steel Square cross-section For 25 mL | 1
85 | Pipette filler/propipette | Silicon rubber bulb | 2
86 | Pipette holders | 5 compartments Wall mounting | 1
87 | Wash bottle | Volume: 500 mL Plastic | 5
88 | Serological pipette | Volume 0.1 mL Made of borosilicate glass | 30
89 | Serological pipette | Volume 1 mL Made of borosilicate glass | 30
90 | Serological pipette | Volume 2 mL Made of borosilicate glass | 30
91 | Serological pipette | Volume 5 mL Made of borosilicate glass | 30
92 | Serological pipette | Volume 10 mL Made of borosilicate glass | 30
93 | Replica plater for nutrient plate | Cylinder made of PVC Ring made of aluminium | 5
94 | Microspoon and spatula weighing set | Double-ended spoons and spatulas-6 different shapes Stainless steel Carrying case | 1
95 | Autoclave indicator tape | Impregnated to show the word “Autoclaved” after 15 minutes of exposure at 120°C Length: 1m | 2
96 | Soil sieves set | Set of 3 sieves Stainless steel Sieve diameter approx 30 cm | 1
97 | Staining jar | Hellendahl pattern | 10
98 | Staining jar | Schiefferdecker pattern | 5
99 | Forceps | Stainless steel Length: 10-15 cm | 5
100 | Forceps, dumont | Made of stainless steel Sharp points Dumont No.5 Length: 10-15 cm | 5
101 | Scissors | Made of stainless steel Straight Length approx: 10 cm | 2
102 | Micro-dissecting scissors | Made of stainless steel Sharp points Straight Length approx: 10 cm | 2
103 | Dry glass bead sterilizer | Temperature: 250 C Chamber depth: 5-10 cm One package of glass beads included | 1
104 | Surber net sampler | Standard type Frame made of aluminium Net made of polyester Mesh size 1 mm Net length approx: 0.5 m Frame dimensions: 0.4 x 0.4 x 0.2 m | 2
105 | Bottom sampler (grab) | Van Veen design Made from stainless steel Surface area: 250 cm² | 1
106 | Aquarium with top | Volume: 200 l Material: glass Thickness: 8 mm Dimensions (l x w x h): 100x50x40 cm | 2
107 | Aquarium lighting | T5 light with 2 x 40 W Tube diameter: 35 mm Length: 850 mm | 2
108 | Eheim air pump | Pump output: 400 l/h Maximum height range: 2 m | 2
109 | Canister filter | For aquariums up to 200 L Pump output: 1500 l/h Filter volume: 5 L | 1
110 | Bench-top autoclave | Chamber volume approx: 80 L 1 shelf | 1
111 | Plant press (professional) |  | 3
112 | Artemia (brine shrimp) hatchery |  | 1
113 | Insect label cards | Dimensions: 3 x 1.25 cm, 1000 pcs | 10
114 | ELISA HIV/AIDS test | Didactic set, includes at least: comboplates, simulated anti-human antibody enzyme linked conjugate, simulated chromagen, simulated patients sera, simulated negative control serum, simulated low positive control serum, simulated high positive control serum, vial with glass beads coated with simulated HIV antigen, Accompanying textbook (description) in English language | 2
115 | Fish skeleton - carp (Cyprinus carpio) | Didactic; complete carp skeleton in natural position | 1
116 | Dove skeleton (Columba palumbus) | Didactic; complete dove skeleton in natural position | 1
117 | Mouse skeleton and stuffed mouse | Didactic; mouse skeleton and stuffed mouse in a natural position; transparent showcase | 1
118 | Fish plastinated slice | Didactic; Plastinated sagital slice of fish, with acrylic protection | 1
119 | Bird plastinated slice | Didactic; Plastinated sagital slice of chicken, with acrylic protection | 1
120 | Mammal plastinated slice | Didactic; Plastinated sagital slice of rat, with acrylic protection | 1
121 | Full-colour dissectograms | Didactic set; dissectograms of at least - Frog - Fish - Clam - Earthworm - Pig Fetus - Grasshopper - Cat | 1
122 | Mitosis model | Didactic set; 3D relief model of animal cell in mitosis | 1
123 | Meiosis model | Didactic set; 3D relief model of animal cell in meiosis | 1
124 | AP Biology Lab 8: Population genetics & evolution | Didactic set; paper pieces with phenythiocarbamide; cards labelled A and a; Accompanying textbook in English language | 1
125 | Polymerase chain reaction didactic set | Didactic set; non-human DNA samples, practice gel loading solution, electrophoresis buffer, Methylene Blue stain, microtipped transfer pipettes; Accompanying textbook in English language | 1
126 | Blood typing with Rhesus factor | Didactic set; dropper bottles of artificial blood (A, B, AB and O), dropper bottles of artificial anti-A, anti-B and anti-Rh serum, washable permanent test trays with 3 wells, Accompanying Information for teacher in English language | 5
127 | Forensic pathology simulation | Didactic set; Simulated Serums (Anti-A, Anti-B, Anti-RH), Blood typing trays, Crime Scene Simulated Neo/BLOOD sample, Suspect Simulated Neo/BLOOD samples, Accompanying description in English language | 1
128 | Simulation of DNA forensics set | Didactic set; Set for simplified DNA fingerprinting; Accompanying textbook in English language | 1
129 | Simulation of DNA paternity testing lab | Didactic set; Set for simplified DNA fingerprinting; Accompanying textbook in English language | 1
130 | Simulation of genetic diagnosis of cancer lab | Didactic set; non-human simulated cancer DNA sample; cancer marker sample; Accompanying textbook in English language | 1
131 | Glass cell model | Didactic; model of human undifferentiated cell (40,000 times enlarged) | 1
132 | Lactose tolerance set | Didactic set; glucose test strips; lactase; Accompanying instructions in English language | 1
133 | Immunology and evolution experimental set | Didactic set; set of simulated blood antigens and antibodies of human and other animals | 1
134 | Complete fermentation system | Didactic set; fermentation vessel noncorrosive thermoplastic headplate, magnetically coupled agitator, inlet air sparger with inlet air filter, exhaust air outlet, tubing assemblies, thermometer, thermometer well, sample tube, air pump, connectors, immersion heater, air flow valve, magnetic stirrer built-in pH meter and probe | 1
135 | Lab Set: Fruitfly genetics | Didactic set for basic fruit fly genetics activities; 15 sheets with snap- apart features of male and female fruit flies | 1
136 | Advanced fossil collection | Set of 30 identified plant and animal fossil taxa; Sheet with descriptive key; Geochronological chart | 1
137 | Pipette pump | For pipettes up to 2 ml Rotating thumbwheel | 5
138 | Pipette pump | For pipettes up to 10 ml Rotating thumbwheel | 5
139 | Pipette pump | For pipettes up to 25 ml Rotating thumbwheel | 5
140 | Terrarium | Made of glass, thick 8 mm Dimensions about 50x50x50 cm | 5
141 | Animal cage | Made of powder coated steel Top opening and front opening doors Dimensions approx: 60 x 35 x 40 cm | 5
142 | Hydroponics complete grow set | Mylar or equivalent Box Dimensions approx: 90x60x160 cm Lamp carrier Lamps MH 400W and HPS 400W Fan Temperature and humidity sensor/tester Timers for triggering lights and fan | 2
143 | Laser leaf area meter | Portable Display: LCD Operating Temperature: up to 50°C Measuring Thickness: up to 15mm Measuring Width: up to 150mm Measuring Length: 35 cm Scanner: 675 nm Laser Diode Scanning resolution: 0.025 mm² | 1
144 | Periodic System display | Large display of pure form periodic system elements including metals, non-metals, weak radioactive elements and gases Each element sample sealed in transparent cube (glass, acryl or similar material) with minimal dimensions of 80x80 mm and engraved name, symbol and atomic number of each element Dimension approx:200cm(W)x100cm(H). Illuminated, engraved wooden frame | 1

## Lot14- Chromatography
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 14: Chromatography
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | Gas Chromatograph with Ion Trap Mass Detector | Gas Chromatograf Column Oven: • Compatible with standard capillary column baskets.• Typical cool-down time: 350°C to 50°C in 270s • Settable temperature ramp: 0.1°C to 120°C Split/splitless injector: • The injector must be able to operate with capillary column • The injector should feature an optimized, modular thermal profile for split and splitless injection, solvent vent and direct injection. Temperature range from -160 deg C up to 450 deg C with possibility of LN2, air or CO2 cooling. • The inlet chamber design has to avoid the contact of the sample withot metal surface • The injector must allow timed closure/opening of the purge line; Injection volume up to 1000 ul. Column compability: Accomodates up to 0.53 mm I.D. capillary columns Autosampler: Under complete data system control autosampler with 8 places Trap Mass detector Mass Range: 10 -1000 amu Resolution: Unit mass resolution maintained over the entire mass range EI-sensitivity: External EI SCAN Sensitivity: 500 fg of octafluoronaphthalene gives S/N 30:1 for ion m/z 272; Internal scan sensitivity: 200 ft of octafluoronaphthalene gives S/N 20:1 for ion m/z 272 External Ionization Ion Trap: • Independent temperature control of transfer line, GC heated zones, and mass spectrometer external source heaters. • All instrument parameters and heated zones (injector, GC oven, transfer line, and ion source) are under data system control. • Provides classical, library-searchable full scan mass spectra • Allows Negative & Positive CI as option External Ionization Ion Source: • Source heater control temperature set points between 125-300oC. • Electron impact (EI) ionization with user definable Elec. Energy • Durable, inert materials source Vacuum: High vacuum system using air-cooled 70 l/sec Detector: • Discrete dynode electron multiplier • Linear dynamic range 104 GC-MS Transfer Line: • Transfer line independently heated and under data system control. • Heater control temperature set points between 100 – 350°C. • Vacuum insulated to maintain stable and consistent temperature across the length of the transfer line. MSn capability: • MSn capability, where n=1-5 • Full Scan MS/MS acquisition. • MSn mode of operation can be combined in the same analytical run with full scan, or segments scan Software description • Complete computer control of parameters for GC, liquid autosampler , and mass spectrometer. • Complete integration of autosampler, gas chromatograph, and mass spectrometer methodsand set-up procedures. • Fully automated data processing programs • The software must operate with Windows XP operating system for true 32-bit multi-tasking that allows the simultaneous generation and printing of reports as an acquisition proceeds without interruption. • Real time data analysis • Complete audit trail for GLP and QC purposes– records on who changed what, when and why. No change can be saved without a reason being input. Once the data have been processed the results file contains the audit trail together with all data and processing methods used to generate the report. This approach helps to easily understand how data were acquired and processed. The audit trail helps to achieve sound GLP experimental and analytical procedures. • Reporting package - Fully automated comprehensive and flexible reporting. • Library Browser – Library search and report program by the National Institute of Standards and Technology (NIST). - Spectra can be passed to the Library Browser from the Qual browser for additional searching and manipulation. - User libraries of mass spectral and structural data can be created and edited - For automated qualitative processing, search spectra can automatically be appended to a user library. Desktop data station with monitor and printer • minimum Intel Pentium Dual Core E5300 2.6GHz • 2 GB DDR2 PC6400 800MHz • integrated graphical and sound card • DVD+/-RW Dual Layer Black • 500GB 7200rpm SATA II • MIDI Tower ATX 450W Black • 22.5“ TFT or LED or LCD monitor • laser printer 3 year warranty Installation and training for use | 1
2 | LC-MS system | Binary pump • flow range 0.001 – 5 mL/min, in 0.001 ml/ min increments • Flow precision: <0.07% RSD • Pressure: Operating range 0 – 600 bar up to 5 mL/min • Composition Precision < 0.2 % SD, at 0.2 and 1 mL/min • Gradient formation: High pressure dual mixing/gradient • Leak sensor: Built in • Vacuum degasser: 4 channels. Autosampler • Pressure Operating range 0 – 600 bar • Injection range 0.1 – 100 μL in 0.1 μL increments Column Compartment: • Peltier termosttated, temperature range 10 degrees below ambient to 100 °C • Temperature stability: ± 0.15 °C • Temperature accuracy: ± 0.8 °C with calibration ± 0.5 °C Detector: • Light source: Deuterium lamp • Wavelength range: 190 – 600 nm • Wavelength accuracy: ± 1 nm • Short term noise:± 0.35 × 10-5 AU, at 230 nm • Drift: 3 × 10-4 AU/hr, at 254 nm • Linearity: >2 AU • Maximum data rate: up to 80 Hz • Flow cells: 10-mm cell path length Mass spectrometer: • Mass Range: m/z 10 - 3000 • Mass Accuracy: +/- 0.13 u within the calibrated mass range in scan mode • Mass Axis Stability: +/- 0.10 u or 100 ppm, at least 12 hours • POS/NEG Switching Speed: Up to 30 msec • Scan Speed: 2500 u/s • SIM Sensitivity: 1pg reserpine, for ion m/z 609.3, 100:1 RMS with multimode ionization source – to demonstrate it at installation; Ion source Multimode ionization soource APCI+ESI • Tuning Completely automated tuning with Pre-made calibrants • System must be delivered with PC, printer and LCMS software ; Software must be capable of displaying, acquiring, storing and manipulating mass spectral information and UVVis information simultaneously. System must have operating software that includes functions for system diagnostics, integrated GLP/cGMP tools. Tests and procedures in the system diagnostics can be used to verify that system performance is acceptable on an ongoing basis. Tracking of the status of system maintenance must be available and system must have on-line logbooks that provide date- and time-stamped records of runs, errors, and maintenance events. • System must be delivered with adequate N2 generator. 5 columns included, to be specified in agreement with the supplier 2 guard columns, to be specified in agreement with the supplier spare part kits, filters, rotor seals, frit sets, for 2 year maintenance (preventative maintenance kits with some suppliers ) at least 2 day practical training for 5 persons on site Training for a larger group (10-20people) in Belgrade Desktop data station with monitor and printer dual core processor > 2.5 GHz speed 4 GB RAM (>666MHz bus speed) DVD+/- RW Dual Layer 1TB SATA HDD Windows 7 64bit 22“ TFT or LCD or LED monitor laser printer, mouse and keyboard 3 year warranty Installation and training for use | 1
3 | Water Deionizer | work based on reverse osmosis principle (RO), optimal capacity up to 200 l/h and water quality up to 1 μS/cm System should contain: 1. Mechanic filter – capsule purifier (minimal flow of 1m3/h) 2. Deferisator (minimal flow 1m3/h) with electronic controller for time programming of rinsing process. 3. Double softener (working flow 2 m3/h) equipped with two bottles and containing microprocessor conducting system. AC Filter (minimal flow 1m3/h): 4. RO module with pre filtration unit (aseptic version) with complete microprocessor based control. 5. Adaptable unit for electro deionisation of water, suitable for final preparation of water systems from RO units with double out unit. 6. Reservoirs – compensation units for storage of ultrafine water (Volume 200 l) 7. Interface containing adapter for local IT network. It should enable control and programming from long distance network terminals as well as software enabling control and conducting of the whole deionisation system. Spare set shuld contain: 1. Filter 10" 20-25 micron (4 pieces) 2. Active carbon filter for Aqua big (4 pieces) 3. Filter 10" 5 micron (4 pieces) 4. RO Membrane for Type 4021 (2 pieces) 5. Complete control chemicals for water hardness (1 piece) 6. Chlorine Comparator (1 piece) 7. Fiting PVC set (size D20-25) with the intervention set in the case of leakage of distribution of purified water (1 piece) 3 year warranty Installation and training for use | 1

## Lot15-ICT-Net Equipment
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 15: Net Equipment
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | Media converter Gigabit Optical | Gigabit Ethernet 1000Mb/s Fiber multi-mode Converter 1 SC fiber optic; 1 RJ45 jack, Min. Distance 500 m AC/DC adapter included | 12
2 | Fiber duplex patch | Fiber duplex patch SC-SC 2m, multi-mode | 20
3 | UPS 10.000VA | On-line with double voltage conversion Real sinus wave output Input voltage 110V to 276V, Automatic bypass Graphics LCD display min. 1500 VA / 1350 W Management software with network support RS-232 port to PC Spare Batteries | 3
4 | Rack UPS 19" | On-line with double voltage conversion Real sinus wave output Input voltage 110V to 276V, Automatic bypass Graphics LCD display Smart – microprocessor controlled Management software with network support RS-232 port to PC Spare Batteries 16 x 12V-9Ah | 20
5 | Wireless AP | 802.11b/g 2,4GHz WLAN outdoor – output power 20 dBm = 100mW (software adjustable) N connector for external antenna (no antenna included) Software AirOS - AP / router / client / bridge Integrated PoE splitter PoE injector AC/DC adapter included Outdoor housing Operating temperature: -20C to +70C | 7
6 | Wireless AP with antenna | 802.11b/g 2,4GHz WLAN outdoor – output power 20 dBm = 100mW (software adjustable) Integrated antenna 8 dBi angle 60 / 60 degrees No external antenna connector Software AirOS - AP / router / client / bridge Integrated PoE splitter PoE injector AC/DC adapter included Outdoor housing Operational temperature: -20C to +70C | 8
7 | USB 2.0 Gigabit Ethernet RJ-45 Adapter | Supports 10 / 100 / 1000 Mbps auto-sensing capability Supports Auto-Negotiation and Auto-MDIX RJ-45 Gigabit Port (straight and cross network cable auto-detection) USB specification 2.0, 1.1 and 1.0 compliant Supports USB full and high speed modes with bus power capability Supports TCP/IP, NetBEUI, IPX/SPX protocols IEEE 802.3, 802.3u, and 802.3ab standards (10Base-T, 100Base-TX, and 1000Base-T) compatible Supports both full-duplex and half-duplex operation in Fast Ethernet Supports Jumbo packet of up to 9KB Supports suspend mode and remote wakeup through link-up or magic packet Equipped with diagnostic LEDs Plug and Play Installation Powered by USB port Supports: Microsoft XP/Vista/7 Mac OS x10.3/10.4 Linux 2.6.14/2.6.9 kernel | 10
8 | Usb wireless card | Interface USB 2.0 Antenna Type On-Board Wireless Standards IEEE 802.11n, IEEE 802.11g, IEEE 802.11b Frequency 2.400-2.4835GHz Wireless Security Support 64/128 bit WEP, WPA-PSK/WPA2-PSK | 5
9 | Ethernet Tester | 18 remote units for larger installation Compact in size so everything fits into one pouch Tests for wiring faults of open, short, reversed and crossed pairs Verification of shield integrity Specially designed remote unit that enables one person to test installed cables Built-in Portfinder function to trace ports back to switch Easy to read fault display and high speed testing Low battery indicator Tone Generator Features | 2
10 | Cable Tie Gun | Works with plastic cable ties up to .25" (6.35 mm) in width Neatly fastens & cuts in one action Adjustable tension for changing cable bundle pressure Automatic cable tie cut-off | 2
11 | KVM switch | 4-port Integrated 3-Port USB 2.0 Hub Universal use with USB and PS/2 devices PC selection via push buttons, hot key, or OSD utility Hot plug and play Emulates a keyboard and a mouse Incl. two cable sets with 1,8m length OS independent | 4
12 | Network IP Tester | Network Types: 10/100 Auto Negotiating 10baseT 100baseT 1000baseTX Power Supply AC adapter for extended operation Connector Type 1 RJ-45 (Combined port for LAN or NIC) Environment Operating Temperature 0°C to 50°C Storage Temperature -10°C to 55°C Detection of Speed and Duplex of 10/100/1000 base TX Links Verification of IP Address Assignment with the DHCP Client Test Tests Connectivity Accross a LAN/WAN To Measure Packet Response Time for Network Performance Determination To Captures MAC Address of an IP Address To Transmit Link Activation Signals for Port Identification | 2
13 | NTP Time Server with GPS Reference Clock | 6 channel GPS C/A-code receiver Network Connection RJ-45, 10/100 MBit Network protocols OSI Layer 7: TELNET, FTP, SSH (incl. SFTP, SCP), HTTP, HTTPS, SYSLOG, SNMP Network Time Protocol (NTP): NTP v2 (RFC 1119), NTP v3 (RFC 1305), NTP v4 (no RFC), SNTP v3 (RFC 1769), SNTP v4 (RFC 2030), MD5 Authentication and Autokey Key Management Protocols: Time Protocol (RFC 868) and Daytime Protocol (RFC 867) and Synchronization of IEC 61850 compliant devices GPS antenna incl. converter unit, 20m GPS antenna cable | 1

## Lot16 -Personal IT Equipment
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 16: Personal IT Equipment
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | All-in-one MultiTouch PC | All-in-One Multi-Touch PC 23" Intel i5 CPU (multi-core) 4GB RAM 1 x HDD Integrated web camera Wireless keyboard/mouse Wireless 802.11 b/g/n Windows 7 64bit OS | 3
2 | Computer Nettop | Computer Nettop Min: Intel Atom D27xx CPU 4GB RAM 160GB 2.5" HHD (Hybryd Hard Drive) or 120 GB SSD mini ITX/custom form factor case keyboard, optical mouse no OS required | 75
3 | Computer Netbook | TFT 10.1” display 1024x600 Intel DualCore Atom N570 or AMD DualCore C60 CPU 2GB DDR III RAM 160GB SATA II HDD LAN 10/100/1000, Wireless 802.11b/g/n, HDMI No OS required | 2
4 | Computer Mobile Workstation | Computer Mobile Workstation - laptop Minimum: Intel Core i7-2860QM Processor (2.5 GHz ) Mobile Intel HM65 Chipset with vPro Technology 17.0-inch diagonal LED-backlit WUXGA WVA anti-glare (1920 x 1200) Graphics with 1 GB dedicated GDDR3 video memory 8 GB 1333 MHz DDR3 SDRAM 750 GB 7200 rpm 2.5-inch hard drive Integrated 2 MP Webcam DVD+/-RW SuperMulti with Double Layer LightScribe Drive Bluetooth Wireless LAN 802.11 b/g/n Dualpoint (touchpad plus pointstick) and spill-resistant keyboard with numeric keypad and drain Fingerprint Sensor Spare Battery Windows 7 Professional 64 bit with system recovery DVD Microsoft® Office Pro 2010 Bag | 1
5 | Computer Server Database | Xeon X34xx Server 1 CPU 16GB RAM 2 x 1 TB HDD Rack-mount 1U | 1
6 | Computer Server Domain | Xeon 5xxx Server 2 CPU 16GB RAM DDR3 (max 128) 1TB HDD LAN 2x1Gb Onboard VGA Rack-mount 1U | 2
7 | Computer Server Storage NAS | Network Attached Storage Rack-mount CPU Intel Core2Duo 2.8GHz 2GB RAM 8 HDD Tray Raid 0,1,5,6, JBOD HotSwap Max.cap.16TB 2 Gigabit LAN Redundant HotSwap Power Supply 8 compatible SATA II 1TB, 7200rpm HDD | 2
8 | Developers PC | Desktop PC CPU Intel i7 12GB (3 x 4 GB) Triple Channel DDR3 RAM Triple SLI - 3 x NVidia Tesla Graphics 2 x 1TB HDD Keyboard Laser Mouse 1 MPl Web camera Windows 7 Professional 64bit OS 3 x 23" WideScreen - HD LCD Monitor | 1
9 | External BlueRay R/W | External BlueRay R/W USB & Firewire compatible | 2
10 | External HDD | External hard drive Capacity 1 TB USB powered Anti-shock technology | 5
11 | Personal BACK-UPS | Output Power Capacity min. - 860 W/1500 AV Maximal recharge time: 20 hours Minimum Full Load Run Time - 5 minutes Minimum Half Load Run Time - 15 minutes Cable set included | 35
12 | Scanner | Desktop A4 scanner Optical resolution 1200x2400 USB 2.0 interface 48-bit colour MAC OS X Leopard and Microsoft Windows XP/Vista/Windows 7 support | 4
13 | A3 Book Scanner | Document and book A3 desktop scanner USB A3 scanning speed 2.5 seconds per page @ 300dpi Microsoft Windows XP/Vista/Windows 7 and Mac OS X Leopard support. OCR, image retouching software included | 2
14 | Gamepad | USB Cordless Gamepad with Analog Sticks | 4
15 | PCI TV card | PCI TV card with Composite input | 10
16 | Wall-mount radio clock | Wall-hanged atomic radio controlled clock 45cm diameter with integrated thermometer and moisture meter. | 30
17 | Barcode printer | Bar code label thermal printing device Maximum label width 10 cm (4'') USB connectivity with Microsoft Windows XP/Vista/7 and Mac OS X Leopard support Maintenance during exploitation must not require manufacturer's certified service personnel | 1
18 | Barcode scanner | Handheld barcode reader Minimum speed 300 reading/second Minimum reading angle 42 degrees Minimal reading distance 50 cm With autonomous rechargeable battery supply (NiMH, Li-Ion) | 2
19 | Usb barcode scanner | Light source 660nm visible red LED Depth of scan field min. 25mm Scan speed min. 75 scans/s Interface USB Decoding capabillity China Postal Code, Codabar, Code 11, Code 32, Code 39 (full ASCII), Code 39(Standard), Code 93, Code 128, Delta Code, EAN-8 with add-on code, EAN-13 with add-on code, EAN 128 Interleaved 2 of 5, ISBN / Bookland EAN, ISBT 128, Italian Pharmacy, Label Code IV, Label Code V, MSI / Plessey, Standard 2 of 5, Toshiba Code, UPC-A with add-on code, UPC-E with add-on code User-replaceable cables Auto scan activation | 3
20 | Photocopier BW | A3/A4/A5 BW photocopier ADF (automatic document feeder) Auto duplex support and sorting 20ppm Magnification 25%-400% Network capable (RJ-45) | 1
21 | Photocopier colour | Combo printer/copier/scanner/fax 20 ppm 2-sided copying Ethernet built-in for network printing & PC faxing 33.6 KBPSsuper G3 fax 1000 page memory 500 sheets cassette 100 sheet multi-purpose tray A4 ADF 50 sheets | 2
22 | Printer Laser b/w | A4 USB 18ppm 600x600 dpi 2MB RAM | 20
23 | Printer Laser b/w Network | A4 35ppm 1200x1200 dpi duplex 128MB RAM Up to 3000 monthly duty cycle | 3
24 | Printer Laser colour | A4 30 ppm 600x600dpi 256MB RAM Up to 75.000 monthly duty cycle | 1
25 | Printer Laser colour Network | A4 colour 30ppm 600x600dpi duplex 384MB RAM Up to 75.000 monthly duty cycle | 1
26 | Multifunctional scanner - copier - fax | B/W laser Printing up to: 15 ppm Try: 250 sheets 33.6 Kbps Super G3 fax Hi-Speed USB A4 Drivers for MAC OS X Leopard and Windows XP/Vist/Windows 7 - 32 and 64 bit | 2
27 | Document Sheetfed Scanner | Up to 600 dpi, 48-bit, 256 gray levels Duty cycle minimum 1000 pages (daily) Multi feed detection Document feeder minimum capacity 50 sheets Scan speed (ADF, A4) up to 20 ppm/40 ipm (b&w, color - 200 dpi) Scan size minimum (ADF) 215.8 x 860.6 mm Media Types supported: Paper (plain, inkjet), envelopes, labels, checks, cards (business, plastic ID) Media weight (ADF) from 49 to 120 g/m² Standard connectivity 1 Hi-Speed USB 2.0 TWAIN 2.0 Software included: For Windows: TWAIN driver, Nuance PaperPort, Nuance OmniPage, NewSoft, Presto! BizCard SE; For Mac: Mac Scanning Utility | 2

## Lot17 -Specialised IT Equipment
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 17: Specialised IT Equipment
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | 3D Printer | Monochrome 3D printer (white) Resolution: 300 x 450 dpi Vertical Build Speed: 25 mm/hour Build Size: 200 x 250 x 200 mm Material Options: High Performance Composite, Direct Casting, Elastomeric Layer Thickness: 0.09 - 0.10 mm Number of Jets: 300 (including powder, binder and infiltrant) Formats for Printing: STL, VRML, PLY, 3DS, ZPR, Workstation Compatibility: Windows® 7, Windows® XP Professional and Windows Vista® Business, Compatible software: 3ds, MAX™, Archicad™, AutoCAD™, Rhinoceros™ | 1
2 | Apple AirPort Extreme Base Station | Apple AirPort Extreme Base Station IEEE 202.11 a/b/g/n support for 2.4GHz and 5GHz simultaneously | 1
3 | Apple TimeCapsule | Wireless data backup device Time Machine complaint (implemented in OS Leopard & Bonjur in MS Windows) Minimal capacity 2TB Dual band wireless capable, Airport, Airport Extreme, 802.11 a/b/g/n, support for 2.4GHz and 5GHz simultaneously WPA/WPA2/128bit WE | 1
4 | Handheld GPS | Dynamics 6g Interface USB Integrated quad helix antenna MCX connector for external antenna Display: 3.5 x 5.5 cm Sun reading colour TFT with backlight Water proof 1m depth for 1min with no damage Operating temperature range: -15°C to 70°C Memory autonom | 5
5 | Robotic Arm: Lynxmotion AL5D | 4 DOF (Degrees Of Freedom); Wrist rotate available. 10.25" median reach 13oz lift capacity. All-aluminum construction, SES-compatible.Serial port-based version control PC software. Includ interactive Windows program that uses external inputs to affect the robot's motion in closed-loop projects. Advanced inverse kinematics positioning control using mouse or joystick. | 1
6 | Portable 3d object scanner | Compact design Multi-beam laser scanning - object not rotating while scanning single slice High accuracy/resolution, macro mode up to 0.00005"/400 samples per inch, wide mode 0.015"/150 samples per inch Acquisition speed 50.000 processed points per second | 1
7 | Apple iMac | Minimum: 2.5GHz Quad-Core Intel Core i5 21.5" - 1920 x 1080 resolution 4GB (two 2GB) memory 500GB hard drive AMD Radeon HD 6750M with 512MB 1920x1Magic Mouse + Magic Trackpad Apple Keyboard with Numeric Keypad (English) Mac OS X Snow Leopard 64b | 12
8 | Server Mac Mini | Apple computer - Mac mini with Snow Leopard Server Minimum: 2.66 GHz Intel Core 2 Duo 8 GB RAM 320 GB Serial ATA HD Promise SmartStore DS4600 4x1TB RAID System External USB DVD SuperDrive Apple Mouse Apple Wireless Keyboard Mac OS X Server Snow Leopard | 3
9 | Weather station | Weather station: Indoor Temperature ( 0 °C to + 60 °C, +/- 0.8 °C), Indoor Humidity ( 1 % to 99 % RH, +/- 5%), Outdoor Temperature ( – 29 °C to + 79.9 °C, +/-0.8 °C), Outdoor Humidity ( 1 % to 99 % RH, +/- 5%, Possibility of connecting up to 8 additional transmitters), Dew point display, Windchill display, Atmospheric Pressure (from 300 hPa to 1100 hPa), Wind Speed and Wind Direction ( from 0 to 200 km/h (or from 1 to 49.9 m/s), +/- 0.1 km/h), Wind direction displayed on a 360° compass, Rain ( from 0 to 999.9 mm (over 1 or 24 hours), +/- 0.3 mm). Recording of Min / Max with time and date of readings. Outdoor Transmitter (Wireless Data transmitted at 868 MHz, Wireless Transmission at 868 MHz : min 120 metres (in open field)), PC connection ( with Computer connection, Possible storage of up to 3000 sets of data, Manage information and create graphs) 6 outdoor transmiters | 3
10 | Image acquisition card | Image acquisition card with Digital Camera Link® protocol and acquisition ports compatible with Sony XCL-V500 camera | 1
11 | XMEGA develoment system | Atmel® ATxmega128A1 8Mhz oscilattor Bootloader over usb mikroC PRO for AVR licence | 1
12 | BeagleBoard development kit | BeagleBoard-xM Super-scalar ARM Cortex TM -A8 512-MB LPDDR RAM High-speed USB 2.0 OTG port optionally powers the board On-board four-port high-speed USB 2.0 hub with 10/100 Ethernet DVI-D (digital computer monitors and HDTVs) S-video (TV out) Stereo audi High-capacity microSD slot and 4-GB microSD card JTAG Camera port | 3
13 | Detection System for Library books | Detection system, designed for use with Tattle-Tape or equivalent Security Strips, for protects library items from loss, gates signal with lights or alarm when an item passes through that has not been desensitised, direct mount, base plate, digital patron | 1
14 | Security Strip Application System | Security Strip Application System designed for use with Tattle-Tape or equivalent Security Strips | 1
15 | Detection System - Book check device | Detection system Book check device for desensitize and desensitise nonmagnetic Tattle-Tape or equivalent Security Strips, USB support | 1
16 | IPTV System | IPTV system – 36 channels 3 transponders DVB-S No receiver required 30 HD set-top boxes VOD service | 1
17 | Virtual Reality Head-Mounted Display with dual VGA Input | Portable Powered over USB from PC Integrated Orientation tracker Built-in noise cancelling microphone and headphones. Professional head strap. 800 x 600 resolution, ~40 degree field of view. Software: Import 3D and multimedia resources, Directly connect to VR hardware | 1
18 | Virtual Reality position tracker, InterSense InertiaCube2+ (Europe version) | Sourceless tracking with full 360° range, 180 Hz update rate with adjustable motion prediction, Adjustable output filters and rotational sensitivity, dll for Windows Vista/XP , so for Linux, SDK for Windows and Linux, Software libraries support up to 32 s | 1
19 | Blower | Operation free from static electricity. With dust bags for vacuum mode Variable speed Control dial. Continuous rating Input 600W Air volume 0 - 2.8 m3/min Air pressure 0 - 5.5 kpa No load speed (rpm) 0 -16,000 | 1
20 | Air compressor | 230V/50Hz min: 1800W No cylinders 1 Air flow 250 l/min - 195 l/min Work pressure 8 bars Max prssure 10 bars Capacity 50 l Weight up to 35 kg Oil-free pump Vertical portable-with wheels ASME code compliant Compressor set and hose included Pump and motor located at the bottom of the tank | 1
21 | Power supply tester | ATX PSU each output connectors check Tests 20-pin and 24-pin power supplies Tests SATA power Tests Pentium 4 power connector Tests PCI-Express power connectors Tests Xeon power connectors Tests Floppy drive connectors Tests standard 4-pin power supply connectors Accurate voltage indicator +/-V(+12V1/+5V/+3.3V/5VSB/+12V2/-12V) ATX P.G value display Lower or higher P.G values alarm Lower voltage detected alarm Over voltage alarm No voltage detected alarm | 3
22 | PC cable tester | Connectors: IDC40, HDB15M/F, Audio, RJ11, RJ45, e-SATAp, SATA, Display Port, USB3.0(A), USB3.0 Power B, USB Micro AB, USB Mini-B, HDMI, Mini HDMI, Micro HDMI, IDC40 extension port Control switch: On/Off, up, down, save, delete, back and enter Display: LCD size, 2 rows–16 characters and back light Battery: DC 9V battery compartment, USB interface for power supply | 1
23 | PC Top Deck Tech Station | Two tier work bench Add-In Card support brace with 10 thumb screws 2x HDD Rails for storage/cooling for four 3.5" or six 2.5" hard drives Secure rubberized motherboard standoff system 5x 120mm fans: 2x HDD cooling, 3x mobo/vid card cooling Neoprene mat for non-slip lower bench surface. Assembly & usage details ATX Control Kit: Mini on/off & reset switches Power & HDD LED's System speaker | 1
24 | PCI LCD Diagnostic Test Card | Two LCD displays with full text feedback At least 5 BIOS brand codes External module shows POST codes on separate LCD 10 POST codes in memory | 1
25 | USB 2.0 to IDE/SATA Adapter | USB 2.0 to IDE/SATA Cable Support SATA HDD up to 1000GB(1TB) , IDE HDD up to 750GB. Connect any device with an IDE/SATA Interface to a PC with USB interface Transfer rate Approx. 480 MBps Complies with USB 2.0 standards Backwards compatible to USB 1.1 standards Complies with ATA/ATAPI-66 Spe. V1.0 Supports IDE/SATA and enhanced IDE hard drives Supports ATA/ATAPI/SATA CD-ROM/R/RW DVD-ROM and more External power adapter for IDE/ATAPI/SATA devices Linux, Mac, Windows compatible | 5
26 | Acu screwdriver | Battery type NI-CD Battery weight max. 2.1 kg Rotation speed (n) 0-500 / 0-1400 min-1 Battery capacity min: 1500 mAh Max. torque 23 Nm charger time max: 1.5 h Chuck type 0,8-10mm keyless chuck + auto lock function Material housing ABS soft housing Quick stop True Number of speed settings 2 BMC True Soft grip True LCD indicator True Charger True Batteries 2 x 1,5mAh Spare batteries | 3

## Lot18 -Software
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 18: Software
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | CorelDraw Graphics suite X5 | CorelDraw Graphics suite X5 | 2
2 | Creative Suite 5 Design Standard | Complete include: Adobe Illustrator® CS5, Adobe® InDesign® CS5, Adobe Acrobat® 9 Pro, Adobe Bridge CS5, Adobe Device Central CS5. | 2
3 | IBM SPSS 20 Statistics Standard Edition | IBM SPSS 20 Statistics Standard Edition | 1
4 | MS Office | MS Office Professional Academic 2010 | 30
5 | MS Office | MS Office Pro 2010 | 1
6 | QuarkXPress 9 | QuarkXPress 9 | 1
7 | Software Mathematica | WolframResearch, Mathematica 8, Linux ver. | 1
8 | Software Mathematica | WolframResearch, Mathematica for Students, Linux ver. | 10
9 | Software Matlab | MATLAB & Simulink Student Version Release 2011a - Linux | 10
10 | Software OS MS Windows Win 7, business, N class | Software OS MS Windows Win 7, business, N class | 15
11 | Software ResourceMate® v 3.0 | ResourceMate® v 3.0 Plus + Multi-User + ALL add-ons + Extended Support | 1
12 | The Bat! mail client | Mail client - The Bat! Pro v 5 | 10
13 | Meteor detection software | UFO-Capture | 7
14 | Sound Forge | Sound Forge Pro 10 | 1
15 | ABBYY set | Included: ABBYY FineReader 11 OCR Professional Edition (for Windows) with serbian language OCR ABBYY FineReader 11 Corporate Edition ABBYY PDF Transformer 3.0 FineReader Engine 9.0 CLI for Linux for 12,000 page per year | 1
16 | ESRI | ArcGIS for Desktop - Basic (for Win OS) | 1

## Lot19 -AV & Photo Equipment
Project: Public Sector Research and Development
Sub Component A: Petnica Science Centre
Technical Specifications & Technical Offer
Tender: Lab Equipment |  | Lot 19: Audio & Video Equipment
No. | Item | Technical Specification Requested | Quantity | Technical Specifications Offered | Notes, remarks, ref to documentation | Evaluation Committee’s notes
1 | 5.1 Surround System | 200-watt rms Frequency Response: 40Hz~20KHz Subwoofer: 100W Centre/Front/Rear: 20 W Signal-to-noise ratio: 80dB 3 AUX in Power on/off switching IR remote control 6RCA to 6RCA cable; Length: 1.5 m 3.5 mm to RCA cable; Length: 1.5 m Audio wires suitable for system power; Length: 65 m | 5
2 | 3D projection screen | 3D non-depolarizing Front projection screen Portable Telescope metal case Viewing size: 2000 x 1600 mm | 1
3 | 2.1 HiFi Speaker System | 3-Piece 2.1 HiFi Speaker System: min RMS 30W Subwoofer: 12W Satellite: 9W Frequency Response: 20Hz - 20 kHz Minimal Sinal to Noise Ratio: 60dB | 10
4 | 100 V HiFi Audio Ampliifier | 120 W power audio amplifier 4 inputs 1 amplifier output channel Constant-voltage outputs - 100 V For commercial and industrial use Any input can be sent to any output Priority muting | 3
5 | 5.1-channel stereo headphones with microphone | Response Bandwidth: 50Hz-20KHz Power: 0.02W Power supply - rechargeable AA or AAA batteries | 5
6 | Digital Visualizer | ½" CCD and 1,024 x 768 XGA resolution 12x optical and 4x digital zoom Automatic white balance Auto focus Different image modes including image freeze VGA and USB inputs Separate fluorescent and cold cathode lamps Slide and negative film holders 15 frames per second 32 picture memory AV cables, adapters and full-function remotes PS/2 mouse | 1
7 | DLP Interactive Projector | DLP Interactive Projector, DLP technology by Texas Instruments or equivalent, Interactive Pen, 16:10 Native Aspect Ratio, Native Resolution WXGA, HDTV compatible, 2500-3000 ANSI Luments, AC 220V / 50 Hz Spare Lamp | 2
8 | DVD & VHS Combo player | Standalone DVD & VHS player VHS specs: 4-head. DVD specs: MP3, JPEG, WMA, DiVX, Dolby Digital standard, srt,sub, txt titles support. | 2
9 | DVR | High Definition Digital Video Recorder (HD DVR) Recording directly from cable TV and satellite set top boxes at up to 1080i Records in AVCHD format for burning Blu-ray DVD discs Support for standard definition composite and S-Video inputs PC connectivity | 2
10 | Electric cinema projection screen | Image Diagonal Size: 380 cm Aspect Ratio 4:3 Screen Surface: Matte White Washable | 3
11 | Hi-Fi audio system | 2x30W RMS FM Radio Automatic equalizer Mobile devices inputs USB direct recording Audio CD/CD-R/CD-RW/MP3 support Connectivity: 1xheadphones( mini-phone stereo 3.5 mm ), 1 x audio line-input ( mini-phone stereo 3.5 mm ) | 6
12 | Projection Screen | Movable cinema projection screen Image diagonal size: 200 cm Aspect Ratio 4:3 Screen Surface: Matte White Washable | 3
13 | Projector DLP | Native XGA 2700 ANSI Lumen Lamp lifetime 3000/4000 hours Set of cables Spare Lamp Bag | 4
14 | Projector DLP | Native XGA Compressed UXGA, WSXGA+ 4000 ANSI Lumen 2300:1 contrast Lamp lifetime 2000/3000(ECO) hours Wireless 802.11b/g, 2xDSUB, 1xComposite, 1xS-Video, 1xDVI, 1xUSB, 1xComponent, 2xPC Audio, HDMI (HDPC)x1, RS232, RJ45, IR and wired remote control with l Spare Lamp | 1
15 | Plasma TV | Plasma TV 50" Full HD Analog/DVB-T/C tuner for Serbia broadcast | 13
16 | Plasma 3D TV | Plasma 3D TV 50" Full HD 33 active glasses Analog/DVB-T/C tuner for Serbia broadcast | 1
17 | Wall Mount Bracket for TV | VESA mount bracket Vertical Tilt +/- 15 degrees Supported weight minimum 50 kg | 20
18 | TV Plasma Floor Stand | VESA mount bracket Support for 50" Plasma Supported weight minimum 50kg Adjustable height max height to center of bracket >= 150 cm Tilted bracket | 12
19 | LCD TV | LCD TV 42" Full HD Analog/DVB-T/C tuner for Serbia broadcast | 10
20 | LCD TV | LCD TV 32" Full HD Analog/DVB-T/C tuner for Serbia broadcast | 7
21 | Digital Grand Piano | Digital piano instrument able to serve as independent unit or be connected with computer Natural-wood 88 full-size keyboard 3-level AWM Dynamic Stereo Sampling tone generator 30 banks of sampled sounds Polyphony: 64 notes Voices: Grand Piano 1,2; Electric Audio amplifier minimum 300 W, and two speakers Cable set | 1
22 | Translation conference system | Conferencing/simultaneous translation audio system 50 IR concurrent listening users 5 table top microphones 2 wireless microphones Command interface for translators for up to 3 languages Equipment storage boxes Cable set | 1
23 | Digital voice recorder | Portable USB 2.0 capable Memory capacity 1GB NAND flash Mass storage device compatible MPEG1 layer 3 (MP3) recording support 64kbps sample rate 22KHz Integrated microphone and speaker Connector for external microphone 3.5mm audio out connector Power suppl | 16
24 | DV camera | Handheld camcorder Full HD 1080p resolution Recording on SD/SDHC flash memory 37mm wide angle lens optical image stabilizer 4 megapixel digital still photo LCD Display: 2.5 inches Optical zoom: 10 x Maximum Aperture Range: F/1.8-2.5, Optical Sensor Reso | 5
25 | Camera Backpack | Capacity:one camera and 4–6 lenses Size(Interior): 27 X 15 X 42.5 cm Size(Exterior): 30.3 X 25.3 X 46 cm Outer fabric: 600D Polyester, 600D Polyester Ripstop, and High Flow Mesh Nylon Weight: 1.6kg | 1
26 | Camera Stand | Metal tripod Max Height approx 158 cm Metal head Larger camera platforms which incorporates dual axis bubble levels to achieve level shots Support up to 4 kg Bag | 2
27 | Camera Stand | Plastic tripod Max Height approx: 100 cm Bag | 3
28 | Camera UV filter | Ultraviolet filter type Filter diameter: 67mm | 2
29 | Camera UV filter | Ultraviolet filter type Filter diameter: 58mm | 1
30 | Camera UV filter | Ultraviolet filter type Filter diameter: 52mm | 3
31 | Converter for SLR Camera | Tele converter 1.4x For Cannon SLR Camera | 1
32 | Digital exposure meter for ambient and flash light | Ambient light: Incident light EV 0 to EV 19.9. Reflected light EV 0 to EV 19.9 (with 40° Lumigrid) Flash: Incident light f/1.4 to f/90.9 (approx. f/124). Reflected light f/1.4 to f/90.9 (approx. f/124) with 40° Lumigrid | 1
33 | Flash head | Flash energy: 500 ws Flash Colour Temp approx: 5500K Halogen modelling lamp: 100 W Flash Duration 1/1200 – 1/800 sec | 2
34 | Flash head | Flash energy: 300 ws Flash Colour Temp approx: 5500K Halogen modelling lamp: 100 W Flash Duration 1/1200 – 1/800 sec | 1
35 | Illuminator reflector panel | Silver/White Panel High-carbon steel frames Dimension: 80 cm | 1
36 | Illuminator reflector panel | Silver/White Panel High-carbon steel frames Dimension: 60 cm | 2
37 | Illuminator reflector panel | Silver/White Panel High-carbon steel frames Dimension: 30 cm | 2
38 | LED video light panel | 1200LED panel with U Bracket design 110V-230V power adapter and cord 3x Color Filters 14.8V DC power 5400K LED color temperature with 3 gels Light intensity: 1m:9500 lux, 2m:3300 lux 3m:1600 lux Dimenzion: 40.6 x 30.4 x 5 cm | 3
39 | LED Macro Ring Flash - complet | LED Macro Ring Flash for Cannon SLR Camera: Flash GN: 15 Effective Flash Scope: 5CM - 1.5M Colour Temperature: 5500K Exposure Time: 1/100 seconds Power Controller 7 Lens Adapter Rings for lens Cannon SLR camera: 52mm, 55mm, 58mm, 62mm, 67mm, 72mm, 77mm. | 2
40 | Lens | Focal Length Range: 300mm Aperture Range - Wide / Long f/2.8L Focus Adjustment Inner focusing system with USM For Cannon SLR Camera | 1
41 | Lens | Focal Length Range: 15mm Fisheye Aperture Range - Wide / Long f/2.8 For Cannon SLR Camera | 1
42 | Lens | Focal Length Range: 100mm macro Aperture Range - Wide / Long f/2.8L Focus Adjustment: Inner focusing system with USM For Cannon SLR Camera | 1
43 | Lens | Focal Length Range: 85mm Aperture Range - Wide / Long f/1.8 Focus Adjustment: Rear focusing system with USM For Cannon SLR Camera | 1
44 | Lens | Focal Length Range: 18-135mm Aperture Range - Wide / Long f/3.5-5.6 Focus Adjustment: Inner focusing system For Cannon SLR Camera | 1
45 | Lens | Focal Length Range: 50mm Aperture Range - Wide / Long f/1.4 Focus Adjustment: Overall linear extension system with USM For Cannon SLR Camera | 1
46 | Lens | Focal Length Range: 60 mm Macro Aperture Range - Wide / Long f/2.8 Focus Adjustment: Rear focusing system with USM For Cannon SLR Camera | 1
47 | Light stand | Metal light stand Height range approx: 100-250 cm Load capacity approx: 10 kg Air bag for flash | 6
48 | Light box for photography | Material: Perfect white nylon. Dimensions: 60x60 cm | 1
49 | Memory card | Secure Digital High Capacity (SDHC) Capacity: 4GB | 4
50 | Memory card | Extreme SDHC Class 10 Capacity: 8GB | 4
51 | Mini digital studio kits | Color Temperature 5000K +/- 10% 48W Mini Fluorescent Light (main light) 2 x 32W Mini Fluorescent Light (for fill and overhead light) 48W + 32W Mini Fluorescent Lights (for back and bottom lights) Collapsible Desk-top Shooting Table: approx. 69 x 58 x 55 cm Light Stand 2 Table top light stands | 1
52 | Optical Lens and Digital SLR Cleaning Kit | Brush Microfiber Cloth Fluid & Tissue Lens Pen Manual Blower | 2
53 | Photo camera | 18 MP APS-C CMOS sensor ISO 100-6400, H:12800 Full HD movies iFCL metering Spare battery | 2
54 | Photo camera | 15 MP APS-C CMOS sensor Full HD movies High ISO up to 12800 9-point AF system Integrated Cleaning System Spare battery | 2
55 | Photo camera | Effecitve pixels: 14 MP Image senzor 1/2.3-inch CCD with primary color filter Image stabilizer: CCD - shift type High ISO up to 6400 Focus distance Normal, Macro and Super Macro AVI movies Sound WAVE format | 5
56 | Digital camera | Sensor resolution: min. 10 MP Video recording: 640x480 (letter 4:3) and 848x480 (wide 16:9) 30fps in duration not less than 5min 4x optical zoom ISO 80-1000 Shutter speed 8-1/2000 sec integrated TFT display minimum 2.5''size with resolution at least 230kPixel AV output (with cables) USB connector (with cable) Internal memory approx 50 MB SD or SDHC memory support Power supply: Li-Ion battery included with camera and battery recharger Carrying all-weather protective bag. | 11
57 | Photo camera white GPS | Camera white GPS and comapss Effecitve pixels: 16,2 MP Optical image stabilizator Optical zoom: 10x Video clip | 3
58 | Shooting Table | Dimension approx: 50 x 100 cm Non-reflective opaque surface | 1
59 | Shoulder Bags | Capacity:camera and 3-4 lenses Size(Interior): 25 X 15.5 X 18.5 cm Size(Exterior): 27.5 X 21.5 X 22.5 cm Outer fabric: Water-resistant micro fiber, ripstop nylon and 600D TXP™ Weight: 0.85kg | 2
60 | Square-shaped Soft BOX | Square Dimensions: 60x90 cm Metal holder | 2
61 | Square-shaped Soft BOX | Square Dimensions: 60x60 cm Metal holder | 1
62 | Wireless flash trigger | Battery Triggering delay: 1/1600s Synch Speed: 1/125s(focal plane shutter) Range: 15m | 2
63 | Waterproof camera white GPS | Waterproof camera white GPS, compass and manometer Effecitve pixels: 14.0 MP Image stabilizator: senzor-shift Optical zoom: 5x Video clip | 3
64 | Total station - complet | Head Length: 150 Millimeters Objective Lens Diameter: 45mm (EDM:50mm) Magnification: 30x Field of View: 1°30 Resolve Power: 3 " Min. Focus Distance: 1.3 meters Measurement Range: 1 Prism - 2000 meters Detection: 2 horizontal, 1 vertical Method: Liquid type Compensating Range: ±3' Internal Data memory: 24000 pts Heavy Duty Wood Tripod open (178cm) closed (110cm) Single Tilt Prism Set 62 mm - Orange 2.5m Laser Etched Grad TwissLoc w/Precise Tip TwissLoc Prism Pole Bipod Prism Bgg 3 year warranty | 1